## ToDo
Build OpenAPS knowledge |  | https://docs.google.com/document/d/1Pkt9LOy3BGSdh1vkVvxSxpfDFFI1ZFqUY05R02E7OmU/edit#
Finish reviewing requirements
Level 2 Loss Scenarios
Maybe we should do an analysis on the control action 'Deliver Insulin'
Change names to make them all consistent
Checklist:
X second variables
Couple of Control Actions that may not have been in UCAs
Just needed to sycnhronize and formalize loss scenarios
r

## Losses, Hazards and SCs
Losses | Hazards | Scenario Constraints
L1: Loss of life or cognitive ability (hypoglycemia, hyperglycemia, DKA) | H1: Too much insulin (L1, L3) | S1: OpenAPS should not recommend an insulin dose of more than X mg in any 1 hour (H1)
L2: Loss of or damage to equipment | H2: Too little insulin (L1, L3) | S2: OpenAPS should not recommend an insulin dose of less than X mg in any 1 hour (H2)
L3: High variance of glucose levels | H3: Electrical malfunction (L1, L2, L3) | S3: If OpenAPS recognises that the user has hazardous levels of insulin or glucose, OpenAPS should alert the human operator (H1, H2)
| H4: Too much burden on system (L2) | S4: OpenAPS should not send signals to other components which might make those other components malfunction (H3)
S5: OpenAPS should not send so many signals per hour to other system components such that the other components drain battery too quickly (H4)

## Level 1 UCAs
Controller | Control Action | Not Providing | NP Controller Constraints | Providing | P Controller Constraints | Timing | T Controller Constraints | Duration | D Controller Constraints | Notes
OpenAPS | STB: Set Temporary Basal | UCA-STB-1 : OpenAPS does not set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation | CC-STB-1 : OpenAPS must set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation | UCA-STB-100 : OpenAPS algorithm recommends too high temporary basal dosing recommendation when current glucose levels are low and not likely to rise (bases on active insulin and human insulin pump input) | CC-STB-100 : OpenAPS algorithm must not recommend a too high temporary basal dosing recommendation when current glucose levels are low and not likely to rise (based on active insulin and human insulin pump input) |  |  |  |  | (*1) Insulin pump settings should be: Temp Basal type as units per hour not %Carb ratios to grams, not exchange unitsMax basal rate set to a reasonable value (no more than 3 x 4 times regular basal)Set basal profile, carb ratios, ISF valuesDIA should be 5, 6, 7 not 2 or 3Periods where the pump normally has basal settings of 0
|  | UCA-STB-2: OpenAPS does not set the temporary basal dose when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation | CC-STB-2: OpenAPS must set the temporary basal dose when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation | UCA-STB-101 : OpenAPS algorithm recommends too low temporary basal dosing recommendation when current glucose levels are high and not likely to fall (based on current glucose levels, active insulin and human insulin pump input) | CC-STB-101 : OpenAPS algorithm must not recommend too low temporary basal dosing recommendation when current glucose levels are high and not likely to fall (based on current glucose levels, active insulin and human insulin pump input)
|  |  |  | UCA-STB-102: OpenAPS makes a temporary basal recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting) | CC-STB-102: OpenAPS must not make a temporary basal recommendation when there is inadequate and incomplete data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
|  |  |  | UCA-STB-103 : OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS(*1) | CC-STB-103 : OpenAPS must not provide a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)
|  |  |  | UCA-STB-104 : OpenAPS makes a temporary basal recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing) | CC-STB-104 : OpenAPS must not make a temporary basal recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
| ROD: Request operating data (from the insulin pump) | UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data which is from the last X minutes | CC-ROD-1: OpenAPS must get data from the insulin pump when it has not received data which is from the last X minutes | UCA-ROD-100: OpenAPS requests (and uses) old or missing data from the insulin pump in place of recent data (battery fail option) | CC-ROD-100: OpenAPS must not request (and use) old or missing data from the insulin pump in place of recent data (battery fail option) | UCA-ROD-1000: OpenAPS gets data from the insulin pump < X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation | CC-ROD-1000: OpenAPS must get data from the insulin pump at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
|  |  |  | UCA-ROD-101: OpenAPS requests data from the insulin pump more than X times a minute | CC-ROD-101: OpenAPS must not request data from the insulin pump more than X times a minute
|  |  |  | UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation | CC-ROD-102: OpenAPS must request sufficient data from the insulin pump to make a reliable dosing recommendation
| RG: Request glucose level (from nightscout) | UCA-RG-1: OpenAPS does not get valid data from the NightScout when it has not received data which is from the last X minutes | CC-RG-1: OpenAPS must get valid data from the NightScout when it has not received data which is from the last X minutes | UCA-RG-100: OpenAPS requests (and uses) old or missing glucose data in place of recent glucose data | CC-RG-100: OpenAPS must not request (and use) old or missing glucose data in place of recent glucose data | UCA-RG-1000: OpenAPS gets data from the NightScout < X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation | CC-RG-1000: OpenAPS must get data from NightScout at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
|  |  |  | UCA-RG-101: OpenAPS gets data from the CGM more than X times a minute, if NightScout is not available (consequence: could diminish battery or wear out the CGM) | CC-RG-101: OpenAPS must not get data from the CGM more than X times a minute, if NightScout is not available (consequence: could diminish battery or wear out the CGM)
|  |  |  | UCA-RG-102: OpenAPS requests insufficient data from the NightScout to make a reliable dosing recommendation | CC-RG-102: OpenAPS must request sufficient data from NightScout to make a reliable dosing recommendation
| SIP: Suspend Insulin Pump | UCA-SIP-1: OpenAPS does not suspend insulin dosing, or at least provide an audible alert for the human to suspend insulin dosing, when the human is experiencing or close to experiencing hypoglycemia | CC-SIP-1: OpenAPS must have data coming in from the CGM and OpenAPS algorithm to help with the predicting of hypoglycemia CC-SIP-2: OpenAPS must suspend insulin dosing when the human is experience or close to experiencing hypoglycemia | UCA-SIP-100: OpenAPS suspends insulin dosing when the human is not close to experiencing hypoglycemia | CC-SIP-100: OpenAPS must not suspend insulin dosing by the insulin pump when the human is not close to experiencing hypoglycemia | UCA-SIP-1000: OpenAPS suspends insulin dosing more than X minutes after the human is experiencing or close to experiencing hypoglycemia | CC-SIP-100: OpenAPS must suspend insulin dosing within X minutes after the human is experiencing or close to experiencing hypoglycemia | UCA-SIP-10000: OpenAPS carries on suspending insulin dosing for more than X (~120, maybe higher depending on factors) minutes after the human was experiencing or close to experiencing hypoglycemia, causing a DKA risk | UCA-SIP-10000: The insulin pump should not be suspended for more than X (~120) minutes after the human was experiencing or close to experiencing hypoglycemia,
Human Actor | PTR: Put together OpenAPS rig |  |  | UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely (this needs more research, hardware will help) | CC-PTR-100: Human actor must not set up the OpenAPS rig wrong or insecurely (this needs more research, hardware will help)
|  |  |  | UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM) which makes the device incompatible with OpenAPS | CC-PTR-101: Human actor must not make a software update to another device (eg insulin pump, CGM) which makes the device incompatible with OpenAPS
| S: Install or update system dependencies for OpenAPS | UCA-S-1: Human actor does not install all the system dependencies necessary for OpenAPS | CC-S-1: Human actor must install all the system dependencies necessary for OpenAPS | UCA-S-100: Human actor makes a software update to the software that the OpenAPS rig is using, which makes it incompatible with other device(s) in the system (eg pump, CGM) | CC-S-100: Human actor must not make a software update to the software that the OpenAPS rig is using which makes it incompatible with other device(s) in the system (eg pump, CGM)
|  | UCA-S-2: Human actor does not update the OpenAPS system dependencies when there are available updates which are compatible with the other OpenAPS dependencies | CC-S-2: Human actor must update system dependencies when there are compatible updates available | UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS without a "test" configuration | CC-S-101: Human actor must not install unecessary additional system dependencies for OpenAPS without a "test" configuration
| CO: Setup and configure OpenAPS loop | UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly | CC-CO-1: Human actor must update the configurations of OpenAPS accordingly if there is a change in the OpenAPS loop setup which affects the configurations | UCA-CO-100: Human actor selects an inaccurate option when configuring OpenAPS | CC-CO-100: Human actor must not select an inaccurate option when configuring OpenAPS
|  |  |  | UCA-CO-101: Human actor configures OpenAPS with a bad algorithm instead of oref0 | CC-CO-101: Human actor must not configure OpenAPS with a bad algorithm instead of oref0
| SO: Stop OpenAPS loop | UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS or when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations | CC-SO-1: Human actor must stop the OpenAPS loop when there is a problem with OpenAPS or when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations | UCA-SO-100: Human actor provides Switch off (hold button goes into reset) when the system is expected to be up and running | CC-SO-100: Human actor must not provide Switch off (hold button goes into reset) when the system is expected to be up and running |  |  | UCA-SO-1000: Human actor breaks the Power Off button
| UD: Provide useful data to pump (eg carbs) | UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | CC-UD-1: Human actor must provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing | CC-UD-100: Human actor must not provide erroneous input of data affecting insulin dosing | UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime | CC-UD-1000: Human actor must provide meal data to the insulin pump within X minutes from the actual mealtime
| SB: Set bolus on pump |  |  | UCA-SB-100: Human actor sets a bolus (or basal) which unintentionally permanently overrides OpenAPS temporary basals | CC-SB-100: Human actor must not set a bolus (or basal) which unintentionally permanently overrides OpenAPS temporary basals
| PM: Provide mandatory settings for OpenAPS on pump (eg temp basal type as units per hour not %, basal profile, DIA) |  |  | UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS (Temp Basal rates as units per hour not %, carb ratios to grams not exchange units, max basal rate set too high, basal profile / carb ratios / ISF values not set, DIA too low) | CC-PM-100: Human actor must not configure insulin pump settings in a way that is not suitable for OpenAPS (Temp Basal rates as units per hour not %, carb ratios to grams not exchange units, max basal rate set too high, basal profile / carb ratios / ISF values not set, DIA too low)
| UC: Update CGM / CGM receiver | UCA-UC-1: Human does not update CGM when there are updates available for X weeks which can improve glucose data accuracy or prevent problems |  | UCA-UC-100: Human updates the CGM when OpenAPS rig is trying to pull data from the CGM
UCA-UC-101: Human provides wrong CGM update
Infrastructure Developer | UI: Update Infrastructure | UCA-UI-1: Infrastructure developer does not update the OpenAPS infrastructure when there is a problem with the current infrastructure |  | UCA-UI-100: Infrastructure developor makes a faulty update to the OpenAPS infrastructure
UCA-UI-101: Infrastructure develepor makes an update to the OpenAPS infrastructure which makes OpenAPS incompatible with other components in the system being used (Insulin pump, CGM)
UCA-UI-102: InfDev makes an incomplete update to the OpenAPS infrastructure (compared with what has been planned/promised/documented)
| TI: Test | UCA-T-1: Infrastructure developer does not test the OpenAPS tools and infrastrucure sufficiently |  | UCA-T-100: Infrastructure developer tests of the system are flawed |  | UCA-T-1000: Infrastructure developer tests the OpenAPS system infrastructure after it's deployed
UCA-T-2: InfDev does not consider all HW setups/configurations when running/updating the test cases
UCA-T-3: InfDev does not consider each use case for OpenAPS tools and infrastructure when writing tests
UCA-T-4: InfDev does not provide analysis of failed tests
| IF: Forum | UCA-IF-1: Infrastructure developers are not present on the forum that they provide |  | UCA-IF-100: Infrastructure developers provide a flawed forum (bandwidth issues, software bugs) |  | UCA-IF-100: InfDevs do answer the questions more than X hours after the question has been asked
|  | UCA-IF-2: InfDevs are not answering specific types of questions in the forum (bc they are biased, thought somebody else will answer etc), so that some questions just are forgotten |  | UCA-IF-101: InfDevs give wrong information in the forum
IA: Answers
Algorithm Developer | UA: Update Algorithm | UCA-UA-1: Infrastructure developor does not update the OpenAPS infrastructure when there is a problem with the current infrastructure |  | UCA-UA-100: Infrastructure developor makes a faulty update to the OpenAPS infrastructure
UCA-UA-101: Infrastructure develepor makes an update to the OpenAPS infrastructure which makes OpenAPS incompatible with other components in the system being used (Insulin pump, CGM)
| TA: Test | UCA-TA-1: Infrastructure developer does not test the OpenAPS tools and infrastrucure sufficiently |  | UCA-TA-100: Infrastructure developer tests of the system are flawed |  | UCA-TA-1000: Infrastructure developer tests the OpenAPS system infrastructure after it's deployed
| AF: Forum | UCA-AF-1: Infrastructure developers are not present on the forum that they provide |  | UCA-AF-100: Infrastructure developers provide a flawed forum
AA: Answers
OpenAPS Release | EN: E-mail notifications | UCA-EN-1: OpenAPS Release does not provide e-mail notifications to human operator when there is an update to OpenAPS software or system dependencies available to the human operator |  | UCA-EN-100: OpenAPS Release provides an inaccurate e-mail notification to the human operator |  | UCA-EN-1000: OpenAPS Release provides an e-mail notification more than X days after there is an update to OpenAPS software or system dependencies available to the human operator
|  | UCA-EN-2: OpenAPS Release does not provide e-mail notifications to human operator when a safety or security concern with the OpenAPS system has been found |  | UCA-EN-101: OpenAPS Release provides a repeated e-mail notification to the human operator |  | UCA-EN-1001: OpenAPS Release provides an e-mail notification more than X days after a safety or security concern with the OpenAPS system has been found
| LOS: Latest OpenAPS software | UCA-LOS-1: OpenAPS release does not provide all of the latest algorithm and infrastructure/dependency updates that the developers intended |  | UCA-LOS-100: OpenAPS release provides algorithm updates which are based on infrastructure which is not currently in use |  | UCA-LOS-1000: OpenAPS release provides latest valid updates to human operator more than X days after developer has sent them
UCA-LOS-101: OpenAPS release provides infrastructure which the currently used algorithm is not compatible with
UCA-LOS-102: OpenAPS release provides flawed software
UCA-LOS-103: OpenAPS release does not have a sufficient integrity method (checksums, SHA hash) to check updates
OF: Forum

## Level 1 Loss Scenario Analysis
Controller | Control Action | UCA | True Statement | Unsafe Controller Behaviour Type (Information given to controller is correct) | Controller Belief | Loss Scenario Type 1 | Requirement 1 | Inadequate Information and Feedback Type (information given to controller is incorrect) | Controller Belief | Loss Scenario Type 2 | Requirement 2 | Control Path Scenario - | Loss Scenario 3 | Requirement 3 | Control Path Scenario - | Loss Scenario 4 | Requirement 4 | Control Path Scenario - | Loss Scenario 5 | Requirement 5 | Controlled Process Scenario | Loss Scenario 6 | Requirement 6 | Controlled Process Scenario | Loss Scenario 7 | Requirement 7 | Controlled Process Scenario | Loss Scenario 8 | Requirement 8
OpenAPS | STB: Set Temporary Basal | UCA-STB-1: OpenAPS does not set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation | The insulin dosage is not optimal AND it has been over X minutes since the last temp basal recommendation | OpenAPS does not set the temporary basal dose, though the information received does indicate that the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation. | OpenAPS believes the insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation (but still does not set temporary basal recommendation) | OpenAPS rig does not have adequate data to make a dosing recommendation | The OpenAPS rig must pull ___ data from the other components of the system, every X seconds or whenever this data is available If OpenAPS believes that data is insufficient for a dosing recommendation to be made, OpenAPS should still predict the effect of the fallback dosing recommendation and if this is far out of the targets it should make a dosing recommendation anyway, with a warning | The insulin dosage is not optimal and it has been over X minutes since the last temp basal recommendation, but the information given to OpenAPS does not indicate this | OpenAPS believes that the insulin dosage is good | OpenAPS has received inaccurate glucose data from Nightscout | There should be a way for the human manually to override glucose data from the CGM / Nightscout. OpenAPS should prompt the user every week to calibrate the CGM. | OpenAPS provides set temp basal rate but it is not received by the insulin pump | OpenAPS has gone out of range from the insulin pump and cannot issue a dosing recommendation OpenAPS has had wireless interference from the insulin pump and cannot issue a dosing recommendation | The OpenAPS rig should emit a warning if it's going out of range of the insulin pump If the insulin pump goes out of range of OpenAPS, it should fall back to normal operation | OpenAPS does not provide set temp basal rate but the insulin pump still receives one | Insulin pump falls back to standard operation (not using OpenAPS) | OpenAPS should have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user. | Insulin pump receives the temp basal rate from OpenAPS incorrectly / too late | z N / A (not a Loss Scenario) - signal cannot reach insulin pump early (i.e. before we've sent it). Signal cannot be late (if it is sent the wireless signal will either reach the pump quickly or it will not reach the pump at all) |  | The set temp basal rate command has reached the insulin pump correctly, however the insulin pump does not execute the command | Insulin pump has run out of power | If the Insulin Pump is on low battery, OpenAPS should stop sending it commands and let it fall back to normal operation | Insulin pump doses insulin when no temp basal rate has been set | Insulin pump falls back to standard operation (not using OpenAPS) |  | The set temp basal rate command has reached the insulin pump correctly, however the insulin pump executes the command incoorectly or too late | Liquid in the tubing connector has caused the infusion set to get blocked
[6 empty] | Flawed OpenAPS decision making algorithm (UCA-STB-1: OpenAPS does not set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation) | The OpenAPS algorithm must set a temporary basal dosing recommendation if the insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely |  |  | OpenAPS has gone out of range from Nightscout and not received updated glucose information | If OpenAPS is having problems communicating with other devices in the loop, OpenAPS should indicate this and fallback to the human programmed basal rates OpenAPS shall warn the human actor if it is no longer connected to and/or no longer receiving information from Nightscout. There should be checks for wireless interference every X minutes, and OpenAPS should send a warning to the user if there is a positive result |  | OpenAPS radio stick is not plugged into the OpenAPS rig properly | [8 empty] | Insulin pump uses different units | OpenAPS plausibility check shall include examination of the raw data from the insulin pump to ensure the data is in its expected range OpenAPS must inform the human actor if the insulin pump is set to use the wrong units |  |  |  |  | Insulin pump scheduling causes commands from OpenAPS to be executed too late
[6 empty] | OpenAPS rig hardware has not been set up correctly | If the OpenAPS rig components are not connected together properly there must be an alarm, warning or error The OpenAPS rig components must connect together securely User manual must specify the correct hardware setup Software must check that all required OpenAPS dependencies are present Hardware connections need to be set up to match the configuration the human entered to the openAPS configuration. |  |  | OpenAPS has had wireless interference from Nightscout and not received updated glucose information | If OpenAPS is having problems communicating with other devices in the loop, OpenAPS should indicate this and fallback to the human programmed basal rates | Insulin data is needed X seconds before a dosing recommendation to be incorporated into that dosing recommendation | Environment has noisy radio frequency or electromagnetic field | Communication between components must have an integrity check | [7 empty] | Things stuck in the tube between the infusion set and the insulin pump | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and warn the user if it is not. OpenAPS must alert human actor if the insulin pump has stopped giving feedback OpenAPS should be able to interpret error signals from insulin pump and convey to the human actor
[6 empty] | OpenAPS rig has run out of power | The OpenAPS rig must have as long a battery life as possible The OpenAPS rig must alert the user when on low power The OpenAPS rig must have a safe exit process Docs must have guidelines on avoiding a low power supply |  | OpenAPS believes that it has been less than X minutes since the last temporary basal recommendation | OpenAPS has gone out of range from the insulin pump and not received updated operating information | If OpenAPS is having problems communicating with other devies in the loop, OpenAPS should indicate this and fallback to the human programmed basal rates |  | OpenAPS radio frequency not tuned correctly | OpenAPS wireless network signals must have a radio frequency between X MegaHertz and Y MegaHertz | [7 empty] | Pump physically loses connection to body
[6 empty] | OpenAPS rig dependencies / software has been updated and OpenAPS now lacks compatibiliity with the insulin pump | OpenAPS documentation must direct the human actor to easily use the right OpenAPS software dependencies, perhaps via an automated script If the human actor updates to incompatible software dependencies, there must be a way to restore software dependencies that work for OpenAPS easily When backward compatibility is affected by an update, it needs to be clearly flagged |  |  | OpenAPS has had wireless interference from the insulin pump and not received updated operating information | If OpenAPS is having problems communicating with other devies in the loop, OpenAPS should indicate this and fallback to the human programmed basal rates |  | Equipment failiure | There should be a heartbeat mechanism to monitor activeness between systems | [7 empty] | Insulin pump has run out of insulin | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and warn the user if it is not. OpenAPS must alert human actor if the insulin pump has stopped giving feedback
[6 empty] | OpenAPS insulin dose recommendation algorithm is taking too long | When the OpenAPS system is running, there should not be a period of X seconds with no active dosing recommendation from OpenAPS OpenAPS application shall have a monitoring task that will check if tasks are taking too long to execute and warn the user if this occurs. |  |  | Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | The insulin pump must be set to the correct time. If OpenAPS is using a different time to the insulin pump, OpenAPS should notify the human actor OpenAPS should be able to register human input of glucose levels |  | Insulin pump is blocked physically from the OpenAPS rig by another object | Insulin pump and OpenAPS rig must not be blocked off physically to the extent that they cannot connect, and this should be explained in the docs
[5 empty] | OpenAPS believes that the current insulin dosage is good | Flawed OpenAPS dosing calculation algorithm | The OpenAPS algorithm must set a temporary basal dosing recommendation if the insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation |  |  | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should detect the local timezone and prompt the user to change the time when this changes |  | Extreme environmental changes | OpenAPS has a hardware integrity check and can communicate with all necessary peripherals There should be a heartbeat mechanism to monitor activeness between systems
[6 empty] | OpenAPS processes glucose data from Nightscout incorrectly | OpenAPS glucose level value must match glucose data from the CGM, unless it has been overriden by a human input Processing of glucose data needs to be tested and monitored |  |  | OpenAPS has received inaccurate glucose data from the human to override glucose data from Nightscout |  |  | Explorer board not sitting properly on the Edison | OpenAPS has a hardware integrity check and can communicate with all necessary peripherals
[6 empty] | OpenAPS processes data from the insulin pump incorrectly | OpenAPS glucose level value must match glucose data from the CGM, unless it has been overriden by a human inputProcessing of glucose data needs to be tested and monitored |  |  | The disk for RaspberryPi fills up and prevents the operating system from running and completing its processes
[5 empty] | OpenAPS believes that it has been less than X minutes since the last temporary basal recommendation | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS must verify that it is set to the correct time When changing time zone, OpenAPS must either change to the correct time automatically or prompt the human actor to change the time
[6 empty] | OpenAPS rig is hacked by an imposter | The OpenAPS rig and it's potentially sensitive information packets must be secure and protected from being hacked into
|  | UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation | There is no active temporary basal dose and there is adequate data available to make a justified recommendation | OpenAPS does not set the temporary basal dose, though the information received does indicate that there is no active temporary basal dose and there is adequate data available ot make a justified recommendation | OpenAPS believes the last temporary basal dose has finished and there is adequate data available to make a justified recommendation (but still does not set temporary basal recommendation) | Flawed OpenAPS decision making algorithm (eg after pausing for a breakdown it does not resume issueing doses) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) | The OpenAPS algorithm must set a temporary basal dosing recommendation if the insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | The last temporary basal dose has finished and there is adequate data available ot make a justified recommendation, but the information given to OpenAPS does not indicate this | OpenAPS believes the insulin pump is still using an OpenAPS recommendation | OpenAPS rig has differing timestamps to the insulin pump and believes the last recommendation was made recently
[6 empty] | OpenAPS rig components are not connected together properly | If the OpenAPS rig components are not connected together properly there must be a warning or error The OpenAPS rig components must connect together securely |  | OpenAPS believes that there is not adequate data available to make a new recommendation | OpenAPS has received incorrect data from the human actor which conflicts with the rest of the current data and prevents a recommendation | OpenAPS should make the human verify that the data they enter is correct
[6 empty] | OpenAPS rig has run out of power | If the OpenAPS rig has run out of power the pump must fall back to normal non-OpenAPS operation |  |  | Lack of algorithm programmer understanding of the information level required to make a dosing recommendation
[6 empty] | OpenAPS rig dependencies / software has been updated and OpenAPS now lacks compatibiliity with the insulin pump | OpenAPS must check to see if the dependencies are on a 'compatible list' before allowing the user to make an update to OpenAPS dependencies If there are new updates to OpenAPS dependencies which would be compatible with the rest of the system, they would be added to the compatible list When backward compatibility is affected by an update, it needs to be clearly flagged |  |  | OpenAPS rig has differing timestamps to the insulin pump and believes the last recommendation was made recently | If OpenAPS is having problems communicating with other devices in the loop, OpenAPS should indicate this and fallback to the human programmed basal rates
[6 empty] | OpenAPS insulin dose calculation algorithm taking too long | The time between OpenAPS being sent the data to make the dosing calculation and OpenAPS issuing and insulin dose must be < X seconds OpenAPS application shall have a monitoring task that will check if tasks are taking too long to execute and warn the user if this occurs.
[6 empty] | OpenAPS has not been configured correctly | OpenAPS must be configured accurately
OpenAPS rig is hacked by an imposter
[5 empty] | OpenAPS believes the insulin pump is still using an OpenAPS recommendation | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS must be set to the correct time
[6 empty] | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes its recommendation is still active when informed that it's not) | OpenAPS must be aware of the correct IOB and active dosing (including OpenAPS must be made aware if the insulin pump / has been disconnected from the body or OpenAPS)
[5 empty] | OpenAPS believes that there is not adequate data available to make a new recommendation | Flawed OpenAPS decision making algorithm (either accidently or because the human engineer thinks that more information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) | OpenAPS must have an accurate model of the data necessary to make a recommendation OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely
|  | UCA-STB-100: OpenAPS algorithm recommends too high temporary basal dosing recommendation when current glucose levels are low and not likely to rise (bases on active insulin and human insulin pump input) | Current glucose levels are low and not likely to rise | OpenAPS algorithm recommends too high temporary basal dosing recommendation, though the information received does indicate that current glucose levels are low and not likely to rise (based on glucose levels, active insulin and human insulin pump input) | OpenAPS believes current glucose levels are low and not likely to rise (but still recommends a too high dosing recommendation) | Flawed OpenAPS dosing calculation algorithm | OpenAPS algorithm must use a correct insulin to glucose calculation | Current glucose levels are low and not likely to rise, but the information given to OpenAPS does not indicate this | OpenAPS believes glucose levels are not low | OpenAPS has received inaccurate glucose data from Nightscout
[5 empty] | OpenAPS believes current glucose levels are not too low | OpenAPS is using the wrong glucose data from Nightscout | OpenAPS must use recent glucose level data from Nightscout at all times Nightscout glucose level data must be regularly checked for accuracy by human manual glucose reading OpenAPS shall only use glucose data that has been validated from Nightscout. Nightscout data shall be checked regarding plausibility of its data in comparison with the valid ranges approved by the FDA (e.g. the ones used be the pump and glucose monitor) Nightscout data shall be checked regarding plausibility of state changes Nightscout data shall be checked regarding corruption and should have integrity checks |  | OpenAPS believes glucose levels are likely to rise | Human actor has input too many carbs or gives incorrect information relating to exercise / other factors | OpenAPS should give a warning when human actor has provided abnormal carb data
[5 empty] | OpenAPS believes current glucose levels are likely to rise | OpenAPS does not account for human bolus dosing | OpenAPS must include human bolus dosing in its IOB value |  |  | OpenAPS is not informed of human bolus dosing | Human bolus dosing should be encorparated into the OpenAPS IOB / active insulin figure
[6 empty] | Flawed OpenAPS data analysis / projection of future glucose level | OpenAPS must make a valid prediction of future glucose levels based on available data OpenAPS must take as much useful data as possible for its data analysis OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely |  |  | Insulin pump is configured with the wrong DIA / ISF / carb ratio settings | OpenAPS docs must advise human actor how they can determine the correct settings for DIA / ISF / carb ratio OpenAPS should advise human actor if something seems incorrect with the insulin pump settings
[6 empty] | OpenAPS rig is hacked by an imposter |  |  |  | OpenAPS does not know that the human has missed a meal / unusual circumstances / travelling | OpenAPS should prompt the human for data if glucose levels are X mg lower than expected
|  | UCA-STB-101 : OpenAPS algorithm recommends too low temporary basal dosing recommendation when current glucose levels are high and not likely to fall (bases on current glucose levels, active insulin and human insulin pump input) | Current glucose levels are high and not likely to fall | OpenAPS algorithm recommends too low temporary basal dosing recommendation, though the information received does indicate that current glucose levels are high and not likely to fall (based on glucose levels, active insulin and human insulin pump input) | OpenAPS believes current glucose levels are high and not likely to fall (but still recommends a too low dosing recommendation) | Flawed OpenAPS dosing calculation algorithm | OpenAPS algorithm must use a correct insulin to glucose calculation | Current glucose levels are high and not likely to fall, but the information given to OpenAPS does not indicate this | OpenAPS believes glucose levels are not high | OpenAPS has received inaccurate glucose data from Nightscout
[5 empty] | OpenAPS believes current glucose levels are not too high | OpenAPS is using the wrong glucose data from Nightscout | OpenAPS must use recent glucose level data from Nightscout at all times Nightscout glucose level data must be regularly checked for accuracy by human manual glucose reading OpenAPS shall only use glucose data that has been validated from Nightscout. Nightscout data shall be checked regarding plausibility of its data in comparison with the valid ranges approved by the FDA (e.g. the ones used be the pump and glucose monitor) Nightscout data shall be checked regarding plausibility of state changes Nightscout data shall be checked regarding corruption and should have integrity checks |  | OpenAPS believes glucose levels are likely to fall | Human actor has input too few carbs or gives incorrect information relating to exercise / other factors | OpenAPS should give a warning when human actor has provided abnormal carb data
[5 empty] | OpenAPS believes current glucose levels are likely to fall | Flawed OpenAPS data analysis / projection of future glucose level | OpenAPS must make a valid prediction of future glucose levels based on available data OpenAPS must take as much useful data as possible for its data analysis |  |  | Insulin pump is not operating correctly but OpenAPS is not informed of this | OpenAPS should get feedback from insulin pump about whether the pump has applied the dosing recommendation from OpenAPS
[6 empty] | OpenAPS rig is hacked by an imposter |  |  |  | Insulin pump is configured with the wrong DIA / ISF / carb ratio settings | OpenAPS docs must advise human actor how they can determine the correct settings for DIA / ISF / carb ratio OpenAPS should advise human actor if something seems incorrect with the insulin pump settings
[10 empty] | Human suddenly consumes a lot of food / sugar | OpenAPS should prompt the human for data if the human hasn't entered data in a long time or if glucose levels are X mg higher than expected Human actor must not suddenly consume a lot of food / sugar OpenAPS must warn the human actor if insulin is in low supply
|  | UCA-STB-102: OpenAPS makes a temporary basal recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting) | There is inadequate data to justify the dosing recommendation | OpenAPS makes a temporary basal recommendation though it does have the information that there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting) | OpenAPS believes there is inadequate data to justify the dosing recommendation (but still makes one) |  |  | There is inadequate data to justify the dosing recommendation, but the OpenAPS information does not indicate this | OpenAPS believes that there is adequate information to justify a dosing recommendation | OpenAPS may lack some information that conflicts with information that it has, which means it should hold back on making a decision | OpenAPS should prompt and possibly force glucose monitor information to be consistently manually tested by the human
[5 empty] | OpenAPS believes that there is adequate data to justify the dosing recommendation | OpenAPS is using old / missing / corrupt data | If there is no data or insufficient available, this shold be indicated to the user Validity of used data must be checked. Only correct data shall be used for the openAPS algorithms. |  |  | OpenAPS has not received correct data from the human actor which would conflict with the rest of the current data and prevent a recommendation | OpenAPS should prompt the human with the data it has and get human to verify that it should be applied
[6 empty] | Flawed OpenAPS data analysis / projection of future glucose level | OpenAPS must make a valid prediction of future glucose levels based on available data OpenAPS must take as much useful data as possible for its data analysis |  |  | OpenAPS has received inaccurate glucose data from Nightscout | OpenAPS should prompt and possibly force glucose monitor information to be consistently manually tested by the human
[6 empty] | Flawed OpenAPS decision making algorithm (either a bug or because the human engineer thinks that less information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-102: OpenAPS makes a temporary basal recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)) | OpenAPS must have an accurate model of the data necessary to make a recommendation If the data OpenAPS is programmed to required is received, OpenAPS must make a dosing recommendation OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely |  |  | OpenAPS has received inaccurate glucose data from the human to override glucose data from Nightscout
UCA-STB-103 : Malicious actor hacks OpenAPS system and makes a temporary basal dosing recommendation which overrides the OpenAPS safe requested level in the insulin pump (?)
|  | UCA-STB-103 : OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1) | Insulin pump settings are not configured correctly for OpenAPS | OpenAPS provides a temporary basal dosing rate though it does receive the information that the insulin pump settings are not configurated correctly for OpenAPS (*1) | OpenAPS believes insulin pump settings are configured incorrectly (but still makes a dosing recommendation) | Flawed OpenAPS decision making algorithm (UCA-STB-103: OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) | Insulin pump should not make dosing recommendations if there are incorrect configurations which could ruin OpenAPS recommendations OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | Insulin pump settings are not configured correctly, though OpenAPS is not informed this | OpenAPS does not believe that insulin pump settings are not configured incorrectly | OpenAPS does not pull all of the relevant insulin pump configuration information from the insulin pump | OpenAPS must pull all configurations from Insulin pump
[6 empty] | Insulin pump misconfiguration is not fatal for OpenAPS, and OpenAPS is able to adjust its dosing to cause the correct glucose levels (UCA-STB-103 : OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) | If there are wrong configurations, the human actor should be prompted to change these |  |  | OpenAPS is having problems communicating with insulin pump | If OpenAPS cannot pull configurations from the insulin pump, the pump should go back to its normal operating mode Irregularities in the communication must be detected and actions/safe states used during communication problems
[5 empty] | OpenAPS believes insulin pump settings are configured correctly | Human has not entered some of the optional insulin pump settings and OpenAPS default values for insulin pump settings in preferences.json are not suitable for the user | Documentation must clearly request the human actor how to correctly test and apply all the insulin pump settings which would be necessary to give good dosing recommendations
[6 empty] | OpenAPS does not verify insulin pump configurations | OpenAPS should verify whether insulin pump settings are configured correctly for OpenAPS if it is not human dependent
|  | UCA-STB-104 : OpenAPS makes a temporary basal recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing) | Insulin pump is not in a state to give a dose | OpenAPS makes a temporary basal recommendation though the information received does indicate the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing) | OpenAPS believes that the insulin pump is not in a state to give a dose of insulin (but gives a dosing recommendation anyway) | OpenAPS is making recommendations in case the insulin pump comes back online |  | Insulin pump is not in a state to give a dose, though OpenAPS is not informed of this | OpenAPS believes that OpenAPS is in a state to give a dosing recommendation | There is no feedback from the insulin pump to OpenAPS indicating that the insulin pump is not in a state to give a dosing recommendation | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and warn the user if it is not. OpenAPS must alert human actor if the insulin pump has stopped giving feedback OpenAPS should be able to interpret error signals from insulin pump and convey to the human actor
[6 empty] | Flawed OpenAPS decision making algorithm (eg is not designed to stop making recommendations / readjust its algorithms when the insulin pump is not in a state to give a dose) (UCA-STB-104 : OpenAPS makes a temporary basal recommendation when the insulin pump is not in a state to give a dose) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely |  |  | Insulin pump is not aware that it is not in a state to make a dosing recommendation | OpenAPS should get feedback from insulin pump about whether the pump has applied the dosing recommendation from OpenAPS
[5 empty] | OpenAPS believes that insulin pump is in a state to give a dose of insulin | OpenAPS does not verify whether the insulin pump is online and delivering the insulin doses | OpenAPS should monitor IOB levels and ensure that insulin doses are going through to the pump
[6 empty] | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes that the insulin pump is operating and dosing properly when it is not | OpenAPS must be aware of the correct IOB and active dosing
| ROD: Request Operating Data | UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes | OpenAPS has not received data from the insulin pump which is from the last X minutes | OpenAPS does not get data from the insulin pump though it has not received information of getting insulin data from the last X minutes | OpenAPS believes that it has not received data from the last X minutes (but does not request data anyway) | Flawed OpenAPS decision making algorithm (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS does not have insulin data from the last X minutes, but is informed that it does | OpenAPS believes that it has insulin data from the last X minutes | Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock | OpenAPS requests operating data from the insulin pump but the request is not received by the pump | z Same reasons as for above control action | OpenAPS should only share information to authorised individuals | OpenAPS does not request operating data from the insulin pump, but the insulin pump still receives a request | Insulin pump has received an operating data request from another party (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes ) | OpenAPS shall receive warning from Nightscout if there is an operating request from another party and warn the human actor | The insulin pump receives the operating data request from OpenAPS too late or inadequately | z N/A, see previous control action |  | The operating data request has reached the insulin pump correctly, however the insulin pump does not send its operating data | Insulin pump has run out of power | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and warn the user if it is not. OpenAPS must alert human actor if the insulin pump has stopped giving feedback | The insulin pump sends operating data when no operating data has been requested |  |  | The operating data request has reached the insulin pump correctly, however the insulin pump sends the operating data incorrectly or too late
???? | [5 empty] | OpenAPS only pulls data from the insulin pump when it detects changes on the insulin pump (not a problem, just a different method) |  |  |  | Insulin pump is not aware that it is not in a state to make a dosing recommendation | OpenAPS must have access to a timestamp of when insulin pump data was from, not just when it was sent |  | OpenAPS request is formatted erroneously | OpenAPS data requests must always be comprehensible by the insulin pump Compatible insulin pump software and hardware versions must be specified Correct openAPS request formatting must be specified, verified and monitored | [7 empty] | Insulin software / firmware has been updated and insulin pump now lacks compatibiliity with OpenAPS | OpenAPS must verify that the insulin pump software and firmware is compatible with OpenAPS OpenAPS rig and docs should be able to indicate which insulin pump software versions have been tested When backward compatibility is affected by an update, it needs to be clearly flagged
[5 empty] | OpenAPS believes that it has received data from the last X minutes | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock | [5 empty] | OpenAPS request is designed for a different style of insulin pump or different insulin pump software | OpenAPS must verify that the insulin pump software and firmware is compatible with OpenAPS Compatible Nightscout versions must be specified OpenAPS should give a warning if it is connected to incompatible devices or software versions in these devices
[6 empty] | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | OpenAPS must make checksums and other checks to ensure the data it is using is valid OpenAPS must check the timestamps of when the data it's using is originally from to make sure it's recent data
|  | UCA-ROD-100: OpenAPS requests data from the insulin pump that's missing (battery fail option) | The data from the insulin pump that OpenAPS is requesting is missing | OpenAPS requests data that's missing, though information received from insulin pump does indicate that the data is not available | OpenAPS believes that the data from the insulin pump is missing (but uses it anyway) | Flawed OpenAPS decision making algorithm (UCA-ROD-100: OpenAPS requests data from the insulin pump that's missing (battery fail option)) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS does not get valid recent data, but the information from the insulin pump does not indicate that the data is not valid, recent data | OpenAPS believes that the old/missing insulin pump data that it receives is new data | Insulin pump is not aware that it is not in a state to make a dosing recommendation
[5 empty] | OpenAPS believes that the data from the insulin pump is recent data | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | OpenAPS does not check the data to make sure that it is valid data from the correct time |  |  | Insulin pump sends data to OpenAPS with the current timestamp rather than the time that the data was from | OpenAPS shall have a time expectation of when it receives data from the insulin pump and compare the time expectation with the timestamp - plausibility check
[6 empty] | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock |  |  | Insulin pump sends data without including any indicator that the data is not valid, recent data | OpenAPS plausibility check shall examine raw data received from the insulin pump to ensure that the data has not been corrupted
[10 empty] | Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock
|  | UCA-ROD-101: OpenAPS requests data from the insulin pump more than X times a minute | There is a point at which the number of requests an insulin pump has to respond to is too high, wearing it out or preventing it from doing other things | OpenAPS requests data from the insulin pump too frequently, though it is informed of a request limit | OpenAPS believes that it should not request data from the insulin pump more that X times a minute, but does it anyway | Flawed OpenAPS decision making algorithm (UCA-ROD-101: OpenAPS requests data from the insulin pump too frequently, so that the pump is overloaded with requests (?)) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS should not request from the insulin pump more than X times a minute, but the information does not indicate this | OpenAPS believes that it is not requesting information from the insulin pump too many times | The battery and performance of an insulin pump is not tracked against the use and non-use of OpenAPS
OpenAPS believes that it is making a reasonable number of requests from the insulin pump
|  | UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation | To make a reliable dosing recommendation, ___ data is required from the insulin pump (also depends on how conflicting the data is) | OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation, though the information does indicate that more information is required | OpenAPS believes the information requested is insufficient for a dosing recommendation, but requests it anyway | OpenAPS believes that no other relevant information is available to be requested when there is | OpenAPS must check every X minutes for all the latest information OpenAPS must not use information older than X minutes for dosing recommendation; if OpenAPS has incomplete data, the insulin pump should use the human's programmed basal rates | The data from the insulin pump is insufficient, though the information is that it is sufficient | OpenAPS believes that it has sufficient data from the insulin pump | Lack of algorithm programmer understanding of the information level required to make a dosing recommendation | All code should be reviewed by somebody with the required knowledge before merge
[6 empty] | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the insulin pump (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) |  |  |  | Data is sent from the insulin pump which is actually null data or dummy data | OpenAPS must validate data received from insulin pump
[5 empty] | OpenAPS believes the information requested is sufficient for a dosing recommendation | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely
???? |  | UCA-ROD-1000: OpenAPS gets data from the insulin pump < X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation | Insulin data is needed X seconds before a dosing recommendation to be incorporated into that dosing recommendation | OpenAPS gets data from the insulin pump too late to incorporate the data in the next dosing recommendation, though the information does indicate that it is too late | OpenAPS believes it is too late to incorporate the data into the next dosing recommendation, but gets it anyway | OpenAPS requests information at an independent time to giving dosing recommendations | OpenAPS should use the most recent data any dosing recommendation | It is too late to incorporate the data into the next insulin dose recommendation, but the OpenAPS information is that it is not too late | OpenAPS believes that it has time to incorporate the latest data into the next insulin dose | OpenAPS is not informed of other processes in the schedule | Process scheduling in the OS must ensure enough time for OpenAPS to run its functions
[5 empty] | OpenAPS believes it can incorporate the data into the next dosing recommendation | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | If there have been changes to the OpenAPS algorithm, the implications of any timing changes should be considered
[6 empty] | The time that it takes for OpenAPS to make a dosing recommendation from receiving information has been lengthened, but the OpenAPS algorithm has not been adjusted for this | If the OpenAPS algorithm has a fixed time interval for making recommendations, this time interval should always be updated whenever there is a change in the algorithm's information intake interval or a delay in information intake or processing
[6 empty] | Poor task scheduling from OpenAPS, - operating system, badly implemented call-flow | Correct scheduling must be specified and monitored
[6 empty] | Poor task scheduling from OpenAPS - wrong order algorithm, the dosing recommendation is always calculated before reading data and the recommendation is always made on the base of old data | Correct program flow must be specified and monitored
| RG: Request Glucose Level Data | UCA-RG-1: OpenAPS does not get data from Nightscout when it has not received data from the last X minutes | OpenAPS has not received valid data from Nightscout which is from the last X minutes | OpenAPS does not get valid data from Nightscout though it has not received information of getting glucose data from the last X minutes | OpenAPS believes that it has not received valid data from the last X minutes (but does not request data anyway) | Flawed OpenAPS decision making algorithm (UCA-RG-1: OpenAPS does not get data from Nightscout when it has not received data from the last X minutes) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS does not have valid glucose data from the last X minutes, though it is informed that it does | OpenAPS believes that it has valid glucose data from the last X minutes | CGM clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | CGM clock must be accurate and match RasPi time and real clock time, to nearest X seconds All times should be alligned with NTP or a common time source | OpenAPS requests operating data from the Nightscout but the request is not received by Nightscout | z Same reasons as for above control action |  | OpenAPS does not request operating data from Nightscout, but Nightscout still receives a request | Nightscout has received an operating data request from another party |  | Nightscout receives the operating data request from OpenAPS too late or inadequeately | z N/A, same applies to Nightscout as to the Insulin pump, too early or too late is umpossible |  | The operating data request has reached Nightscout correctly, however Nightscout does not send its operating data | Nightscout is not online | OpenAPS shall warn the human actor if Nightscout is not online | Nightscout sends operating data when no operating data has been requested |  |  | The operating data request has reached Nightscout correctly, however Nightscout sends the operating data incorrectly or too late
[5 empty] | OpenAPS believes that it has received valid data from the last X minutes | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock |  |  | Nightscout sends data to OpenAPS with the current timestamp rather than the time that the data was from | OpenAPS shall have a time expectation of when it receives data from Nightscout and compare the time expectation with the timestamp - plausibility check |  | OpenAPS request is formatted erroneously | Compatible Nightscout versions must be specified Correct openAPS request formatting must be specified, verified and monitored | [7 empty] | Nightscout / CGM software / firmware has been updated and it now lacks compatibiliity with OpenAPS | OpenAPS shall perform compatibility checks with all interfacing applications and software and warn the human actor if there is an incompatibility detected When backward compatibility is affected by an update, it needs to be clearly flagged
[6 empty] | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | OpenAPS shall have time expectation, plausiblity check, and check of timestamp to ensure validity of data |  |  | Lack of glucose data verification | Human actor must verify the CGM readings every X hours with a glucose meter. |  | OpenAPS request is designed for a different version of Nightscout | Compatible Nightscout versions must be specified OpenAPS should give a warning if it is connected to an incompatible version of Nightscout When backward compatibility is affected by an update, it needs to be clearly flagged
OpenAPS has received inaccurate glucose data from Nightscout
|  | UCA-RG-100: OpenAPS requests missing glucose data | The data from Nightscout / CGM is missing (and is treated as new data) | OpenAPS requests data that's missing, though information received from Nightscout does indicate that the data is not available | OpenAPS believes that the data from the Nightscout is missing (but uses it anyway) | Flawed OpenAPS decision making algorithm (UCA-RG-100: OpenAPS requests missing glucose data) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS does not get valid recent data, but the information from Nightscout does not indicate that the data is not valid, recent data | OpenAPS believes that the old/missing Nightscout data that it receives is new data | CGM clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | CGM clock must be accurate and match RasPi time and real clock time, to nearest X seconds All times should be alligned with NTP or a common time source
[5 empty] | OpenAPS believes that the data from the Nightscout is recent data | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | OpenAPS shall have time expectation, plausiblity check, and check of timestamp to ensure validity of data |  |  | CGM sends data to OpenAPS with the current timestamp rather than the time that the data was from | Timestamping needs to match with CGM and RasPi clock and real clock with X seconds Timestamp needs to be accurate time of the reading rather than the time of the data transfer
[6 empty] | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock
|  | UCA-RG-101: OpenAPS gets data from the CGM too frequently, if Nightscout is not available (consequence: could diminish battery of or wear out the CGM) | There is a point at which the number of requests a CGM has to respond to is too high, wearing it out or preventing it from doing other things | OpenAPS requests data from CGM too frequently, though it is informed of a request limit | OpenAPS believes that it should not request data from the CGM more that X times a minute, but does it anyway | Flawed OpenAPS decision making algorithm (UCA-RG-101: OpenAPS gets data from the CGM too frequently, if Nightscout is not available (consequence: could diminish battery of or wear out the CGM)) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely | OpenAPS should not request from the insulin pump more than X times a minute, but the information does not indicate this | OpenAPS believes that it is not requesting information from the CGM too many times | The battery and performance of an insulin pump is not tracked against the use and non-use of OpenAPS | If possible, OpenAPS must monitor insulin pump battery OpenAPS must make less than X requests to the insulin pump per minute
OpenAPS believes that it is making a reasonable number of requests from the CGM
|  | UCA-RG-102: OpenAPS requests insufficient data from the Nightscout to make a reliable dosing recommendation | To make a reliable dosing recommendation, ___ data is required from the Nightscout (also depends on how conflicting the data is) | OpenAPS requests insufficient data from Nightscout to make a reliable dosing recommendation, though the information does indicate that more information is required | OpenAPS believes the information requested is insufficient for a dosing recommendation, but requests it anyway | OpenAPS believes that no other relevant information is available to be requested when there is | OpenAPS must check every X minutes for all the latest information OpenAPS must not use information older than X minutes for dosing recommendation; if OpenAPS has incomplete data, the insulin pump should use the human's programmed basal rates | The data from Nightscout is insufficient, though the information is that it is sufficient | OpenAPS believes that it has sufficient data from the Nightscout | Lack of algorithm programmer understanding of the information level required to make a dosing recommendation
[6 empty] | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the Nightscout |  |  |  | Data is sent from the insulin pump which is actually null data or dummy data | OpenAPS must validate data received from insulin pump
[5 empty] | OpenAPS believes the information requested is sufficient for a dosing recommendation | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-RG-102: OpenAPS requests insufficient data from the Nightscout to make a reliable dosing recommendation) | OpenAPS algorithm should be tested, validated in realtime for accuracy and designed securely
???? |  | UCA-RG-1000: OpenAPS gets glucose data too late to incorporate the data in the next dosing recommendation | Glucose data is needed X seconds before a dosing recommendation to be incorporated into that dosing recommendation | OpenAPS gets data from Nightscout too late to incorporate the data in the next dosing recommendation, though the information does indicate that it is too late | OpenAPS believes it is too late to incorporate the data into the next dosing recommendation, but gets it anyway | OpenAPS requests information independently of giving dosing recommendations | Correct openAPS request timing must be specified, verified and monitored | It is too late to incorporate the data into the next insulin dose recommendation, but the OpenAPS information is that it is not too late | OpenAPS believes that it has time to incorporate the latest data into the next insulin dose | OpenAPS is not informed of other processes in the schedule | Process scheduling in the OS must ensure enough time for OpenAPS to run its functions
[5 empty] | OpenAPS believes it can incorporate the data into the next dosing recommendation | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from rig system clock
[6 empty] | The time that it takes for OpenAPS to make a dosing recommendation from receiving information has been lengthened, but the OpenAPS algorithm has not been adjusted for this | If the OpenAPS algorithm has a fixed time interval for making recommendations, this time interval should always be updated whenever there is a change in the algorithm's information intake interval or a delay in information intake or processing
Poor task scheduling from OpenAPS, - operating system, badly implemented call-flow
Poor task scheduling from OpenAPS - wrong order algorithm, the dosing recommendation is always calculated before reading data and the recommendation is always made on the base of old data
| SIP: Suspend Insulin Pump | UCA-SIP-1: OpenAPS does not suspend insulin dosing, or at least provide an audible alert for the human to suspend insulin dosing, when the human is experiencing or close to experiencing hypoglycemia |  | OpenAPS does not suspend insulin dosing even though OpenAPS recognises that the human is experiencing or close to experiencing hypoglycemia |  | The OpenAPS algorithm never suspends the insulin pump, instead falling back to standard non-OpenAPS operation when it is not working as intended | When the human is close to experiencing hypoglycemia, OpenAPS must suspend the insulin pump or at least audibly alert the human user to suspend the insulin pump, within X minutes | OpenAPS information does not indicate that the human is close to experiencing hypoglycemia |  | CGM / algorithm is giving the wrong reading for the current glucose level | Human actor must verify the CGM readings every X hours with a glucose meter. Algorithm values for glucose levels must be accurate. OpenAPS must notify the human when OpenAPS is suspending insulin dosing.
[6 empty] | The human turns the pump on again and OpenAPS is only configured to suspend the pump once | If the human turns the pump on again after OpenAPS has suspended the pump, if the human still has hypoglycemia then OpenAPS must again suspend the pump, within X minutes of the pump being turned on again |  |  | OpenAPS wrongly projects that the human is not heading towards hypoglyclemia | OpenAPS glucose level projections must be verified, and projection calculations must be refined if projections are regularly inaccurate
|  | UCA-SIP-100: OpenAPS suspends insulin dosing when the human is not close to experiencing hypoglycemia |  | OpenAPS suspends insulin dosing when it OpenAPS recognises that the human is not close to experiencing hypoglycemia |  | Bug in the OpenAPS algorithm | OpenAPS should not suspend insulin dosing when the human is not close to experiencing hypoglycemia | OpenAPS information indicates that the human is close to experiencing hypoglycemia |  | CGM / algorithm is giving the wrong reading for the current glucose level | Human actor must verify the CGM readings every X hours with a glucose meter. Algorithm values for glucose levels must be accurate. OpenAPS must notify the human when OpenAPS is suspending insulin dosing.
[6 empty] | OpenAPS algorithm suspends the insulin pump when there is some problem with the insulin pump itself | OpenAPS should inform the human actor when it suspends the insulin pump (due to a problem with the insulin pump) |  |  | OpenAPS wrongly calculates that the human is heading towards hypoglyclemia | OpenAPS glucose level projections must be verified, and projection calculations must be refined if projections are regularly inaccurate
|  | UCA-SIP-1000: OpenAPS suspends insulin dosing more than X minutes after the human is experiencing or close to experiencing hypoglycemia |  | OpenAPS suspends insulin dosing more than X minutes after the human is experiencing or close to experiencing hypoglycemia when it was informed that the human was experiencing or close to experiencing hypoglycemia X minutes ago |  | Kernel scheduler issue | OpenAPS shall employ a monitoring task to ensure that all safety-related threads in the application are triggered by the kernel on a timely basis (e.g. watchdog timer) | OpenAPS information indicates that the human was close to experiencing hypoglycemia less than X minutes ago |  | OpenAPS or insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | OpenAPS should notify user when other components' clock time differs from RasPi system clock
[6 empty] | OpenAPS algorithm takes too long | OpenAPS must suspend insulin dosing within X minutes after the human is experiencing or close to experiencing hypoglycemia (same as controller constraint)
Processor hanging / Device overheating / systemic problems / low voltage / battery (these should be Covered)
|  | UCA-SIP-10000: OpenAPS carries on suspending insulin dosing for more than X (~120, maybe higher depending on factors) minutes after the human was experiencing or close to experiencing hypoglycemia, causing a DKA risk |  | OpenAPS carries on suspending insulin dosing for more than X (~120, maybe higher depending on factors) minutes after the human was experiencing or close to experiencing hypoglycemia, causing a DKA risk, when it is aware that it should not suspend insulin dosing for as long as X minutes |  | OpenAPS turns off after suspending insulin pumping |  | OpenAPS information does not indicate that it should not suspend insulin dosing for as long as X minutes |  | OpenAPS algorithm is not programmed to resume insulin pumping
Bug in OpenAPS algorithm means it does not resume insulin pumping
OpenAPS disconnects from the network after suspending insulin pumping
Human Actor | PTR: Put together OpenAPS rig | UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely (this needs more research, hardware will help) | The OpenAPS rig must be set up in a specific way | Human actor sets up the OpenAPS rig wrong or insecurely, though is informed not to do this | The human actor does not believe the OpenAPS rig is set up properly, but sets it up like this anyway | Human error (accident) UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely | OpenAPS rig setup instructions should be clear, complete, easy to find and cover all possible hardware | The human actor is not informed of how to set up the OpenAPS rig correctly | The human actor believes that they have set up the OpenAPS rig correctly | Human - Error in OpenAPS instructions (Wrong version number of the software) | There should be a way for errors in the OpenAPS instructions to be reported
[6 empty] | Human - Human actor is experimenting |  |  |  | Human - Incomplete instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[10 empty] | Human - Flaw in the OpenAPS rig hardware setup design which the creators don't realise | OpenAPS rig hardware setup should be tested for robustness and fragiliity
[5 empty] | The human actor believes the OpenAPS rig is set up properly | Human - Human actor is rushing or thinks that they already know how to set up the rig | OpenAPS should send a warning to human actor when human actor hasn't set up the rig matching with the configurations |  |  | Human - Docs out of date (UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely) | Whenever a change is made to the OpenAPS system or the compatible versions change, the docs should be updated
[6 empty] | Human - Human actor reads old / deprecated instructions | There should only be one set of official OpenAPS documentation available for the human actor to read, which contain all information that is in any other instructions found online If there is a change to the system, the documentation should be updated to reflect this |  |  | Human - OpenAPS gives the wrong information to the human (possible wrong version number, possible multiple instances of same program etc.) | OpenAPS docs regarding other components should be reviewed regularly, in keeping with the latest recommended approach
[6 empty] | Human - Human does not see the SD card information in the documentation and uses the wrong SD card | OpenAPS should notify user when SD card is of the wrong size or format Important information in docs should be easy to find in the hardware setup area of the docs |  |  | Human - Human actor cannot find information | The OpenAPS docs must be easy to navigate, with a search bar
|  | UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS | The new device (insulin pump, CGM) software is incompatible with OpenAPS | Human actor makes an update to another device (insulin pump, CGM), which makes it incompatible with the OpenAPS rig, though is informed not to do this | The human actor believes the new insulin pump software is incompatible with OpenAPS but uses it anyway | Human error (accident) UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS | OpenAPS should detect if incompatible insulin pump software is being used and indicate this to the user When new updates have been tested to be compatible, OpenAPS should issue an email notification OpenAPS should list compatible versions in the documentation | The insulin pump software update makes the pump incompatible with the OpenAPS, but the human actor is not informed of this | The human actor believes that the new software on the insulin pump is compatible with the OpenAPS rig | Human - Error in OpenAPS instructions | There should be a way for errors in the OpenAPS instructions to be reported
[5 empty] | The human actor does not think the software is incompatible with the OpenAPS rig | Human - Human misreads compatible version numbers | OpenAPS should check that human is using compatible versions of components. New versions of components should be tested |  |  | Human - Incomplete instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[6 empty] | Human - Human actor reads old / deprecated instructions | There should only be one set of official OpenAPS documentation available for the human actor to read, which contain all information that is in any other instructions found online If there is a change to the system, the documentation should be updated to reflect this |  |  | Human - New insulin pump software has not been properly tested with OpenAPS | Add configuration check between systems including of their software/firmware versions used and only allow system to work with insulin pumps that have been properly tested
[6 empty] | Human - Human thinks that they already know how to set up the rig and don't check instructions (UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS) | Add prompts in OpenAPS UI that insists that human reviews and understands instructions prior to/during/after setting up the rig. Nice to have is an educational video on how to set up the rig |  |  | Human - OpenAPS gives the wrong information to the human (possible wrong version number, possible multiple instances of same program etc.) | OpenAPS docs regarding other components should be reviewed regularly, in keeping with the latest recommended approach
[10 empty] | Human - Docs out of date (when software has been updated) | Any errors that are found in the insulin pump instructions should be listed in the OpenAPS documentation There should be a way for errors in the insulin pump instructions to be directed to OpenAPS
[10 empty] | Human - Human actor cannot find information | The OpenAPS docs must be easy to navigate, with a search bar
| S: Install or update system dependencies for OpenAPS | UCA-S-1: Human actor does not install all the system dependencies necessary for OpenAPS | OpenAPS is designed to use a certain set of dependencies (which can be found in Docs) | Human actor does not install all of the system dependencies, though is informed to do this | The human actor believes all the dependencies are necessary for OpenAPS but doesn't install them anyway | Human error (accident) (UCA-S-1: Human actor does not install all the system dependencies necessary for OpenAPS) | OpenAPS should check that all the necessary system dependencies are installed before looping | The system dependencies are necessary for OpenAPS, but the human actor is not informed of this | The human actor does not believe dependencies are necessary for OpenAPS | Human - Human actor cannot find information (installing dependencies) | The OpenAPS docs must be easy to navigate, with a search bar
[5 empty] | The human actor believes dependencies are not necessary for OpenAPS | Human - Human actor reads old / deprecated instructions | There should only be one set of official OpenAPS documentation available for the human actor to read, which contain all information that is in any other instructions found online If there is a change to the system, the documentation should be updated to reflect this |  |  | Human - Docs out of date (Dependencies script out of date) | Whenever a change is made to the OpenAPS system or the compatible versions change, the docs should be updated
[10 empty] | Human - Bug in dependencies script | The dependencies script must install all of the dependencies correctly
[10 empty] | Human - Vulnerability with one of the dependencies / software that OpenAPS people are unaware of | OpenAPS should use only well trusted and well tested dependencies
|  | UCA-S-2: Human actor does not update the OpenAPS system dependencies when there are available updates which are compatible with the other OpenAPS dependencies | There are updates to OpenAPS dependencies available that are compatible with OpenAPS | Human actor does not update OpenAPS system dependencies when there are compatible updates available, though is informed they can make the updates | The human actor believes there are updates to OpenAPS dependencies compatible with OpenAPS available but doesn't install them anyway | Human error (accident) (UCA-S-2: Human actor does not update OpenAPS system dependencies when there are compatible updates available) | When new updates have been tested to be compatible, OpenAPS should issue an email notification | There are updates to OpenAPS system dependencies compatible with OpenAPS available, though the human is not informed of this | The human actor does not believe there are updates to OpenAPS dependencies compatible with OpenAPS available | Human - Human is not notified when there are updates to OpenAPS are available | There should be e-mail notifications of available OpenAPS updates
[5 empty] | The human actor does not believe that there are updates to OpenAPS dependencies compatible with OpenAPS available | Human - Human does not check their notifications / e-mail | Human must be able to select their preferred frequencies of notifications |  |  | Human - Nobody at all is checking whether there are updates to OpenAPS system dependencies available | The versions of OpenAPS dependencies that are installed should be fixed Somebody should keep a look out and test for major updates to OpenAPS dependencies
[10 empty] | Human - Nobody monitors and tests the latest updates to determine if they are compatible with OpenAPS | Major updates to OpenAPS should be tested
[10 empty] | Human - Human is not prompted to update system dependencies after the initial install (?) | There should be an e-mail or OpenAPS prompt to update dependencies when newer versions are being used
|  | UCA-S-100: Human actor makes a software update to the software that OpenAPS is using, which makes it incompatible with other device(s) in the system (eg pump, CGM) | The new OpenAPS software would be incompatible with other device(s) in the system | Human actor makes a software update to the OpenAPS rig, which makes it incompatible with another component(s) in the system, though is informed not to do this | The human actor believes the new OpenAPS software is incompatible with other system component(s) but uses it anyway | Human error (accident) UCA-S-100: Human actor makes a software update to the software that OpenAPS is using, which makes it incompatible with other device(s) in the system (eg pump, CGM) | OpenAPS should detect if incompatible insulin pump software is being used and indicate this to the user When new updates have been tested to be compatible, OpenAPS should issue an email notification OpenAPS should list compatible versions in the documentation | A software update on the OpenAPS rig would make OpenAPS incompatible with other system component(s), but the human actor is not informed of this | The human believes that a software update on the OpenAPS rig is still compatible with the other system component(s) | Human - Error in OpenAPS instructions | There should be a way for errors in the OpenAPS instructions to be reported
[5 empty] | The human actor does not think that the OpenAPS software is incompatible with other system components | Human - Human error (Human actor defaults to installing the latest software and dependencies and doesn't consider that this could cause a failiure) | OpenAPS docs should advise the human to only install dependencies via dependencies script (or at least see the versions here) and not manually |  |  | Human - Incomplete instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[6 empty] | Human - Human actor reads old / deprecated instructions | There should only be one set of official OpenAPS documentation available for the human actor to read, which contain all information that is in any other instructions found online If there is a change to the system, the documentation should be updated to reflect this |  |  | Human - New OpenAPS software has not been properly tested with all the other devices (insulin pumps, CGM) | Add configuration check between systems including of their software/firmware versions used and only allow system to work with insulin pumps that have been properly tested
[6 empty] | Human - Human actor does not read the instructions about which software versions work | Human actor must get a notification e-mail when sofware versions change |  |  | Human - Human actor cannot find information (working version) | The OpenAPS docs must be easy to navigate, with a search bar OpenAPS should not provide dosing recommendations if it has not been tested to be safely compatible with the software versions of other components in the system
[10 empty] | Human - Docs out of date (working versions information) | Any errors that are found in the insulin pump instructions should be listed in the OpenAPS documentation There should be a way for errors in the insulin pump instructions to be directed to OpenAPS
|  | UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS without a "test" configuration | OpenAPS is designed to use a certain set of dependencies (which can be found in Docs) | Human actor installs the additional system dependencies for OpenAPS, though is informed not to do this | The human actor does not believe that the additional dependencies are helpful, but installs them anyway | Human error (accident) UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS | Human actor should be discouraged from installing any other dependencies that are not in the installation script | The additional dependencies are not helpful to OpenAPS, but the human actor is not informed of this | The human actor believes additional dependencies are helpful to OpenAPS | Human - Human actor cannot find information (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | The OpenAPS docs must be easy to navigate, with a search bar
[6 empty] | Human - Human actor is experimenting |  |  |  | Human - Incomplete docs | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[5 empty] | The human actor believes that the additional dependencies are helpful | Human - Human misunderstanding (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | Covered by Controller Constraint (Human actor must not install unecessary additional dependencies for OpenAPS without a "test" configuration) |  |  | Human - Docs out of date (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | Any errors that are found in the insulin pump instructions should be listed in the OpenAPS documentation There should be a way for errors in the insulin pump instructions to be directed to OpenAPS
| CO: Setup OpenAPS loop and configure settings | UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly | There have been changes in the OpenAPS loop setup which affect the OpenAPS configurations | Human actor does not update the configurations of OpenAPS to match to the new setup, though is informed that they should do this | Human actor believes that the changes in the OpenAPS setup should be reflected in the configurations, but still doesn't update the configurations | Human error (accident) UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly | Where possible, OpenAPS should detect a change in the loop setup and request that the human actor updates the configurations | The changes in the OpenAPS loop setup should trigger a human update to the OpenAPS configurations, but the human is not informed of this | The human does not believe it is necessary to update the OpenAPS configurations after changing the loop setup | Human - Human actor cannot find information (UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly) | The OpenAPS docs must be easy to navigate, with a search bar
[5 empty] | Human actor does not believe that the changes in the OpenAPS setup should affect the configurations | Human - Human forgets about the contents of the configurations or that they need to update the configs at all | If there is a change in the components that OpenAPS is physically connected to, OpenAPS should prompt the human actor to review the configurations |  |  | Human - Incomplete or inaccurate docs, missing out a configuration setting | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[6 empty] | Human - Human does not read docs | OpenAPS safe defaults should be verified as part of each release |  |  | Human - There's a new potential device that hasn't been included in the configs | Developers should try and involve any new device which would improve the functionality of OpenAPS
[10 empty] | Human - Vulnerability to malicious actor interacting | OpenAPS should follow good secure coding practices OpenAPS rig should ask for verification when the human actor has requested information
|  | UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS | OpenAPS should be configured a certain way based on the system setup and user's preferences | Human actor selects the wrong option when configuring OpenAPS, though is given all the necessary information to select their preferred option | The human actor understands all the options but still selects the wrong option | Human error (accident) UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS | OpenAPS should make the human verify their selections after configuring OpenAPS | The human is not given all of the right information about the system setup to select the correct OpenAPS configuration | The human believes the system is set up differently to how it is | Human - Human actor cannot find information (UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS) | The OpenAPS docs must be easy to navigate, with a search bar
[5 empty] | The human actor does not understand all the options | Human - Human does not read docs | OpenAPS safe defaults should be verified as part of each release |  |  | Human - Incomplete or inaccurate docs, missing out a configuration setting | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[10 empty] | Human - Vulnerability to malicious actor interacting | OpenAPS should follow good secure coding practices
[10 empty] | Human - Error in OpenAPS instructions | There should be a way for errors in the OpenAPS instructions to be reported
|  | UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0 | The algorithm that the OpenAPS uses should be the latest oref0 | Human actor configures OpenAPS with an algorithm an algorithm other than the latest oref0, though is informed not to do this | The human actor knows that they should use the latest oref0 but uses a different algorithm anyway | Human error (accident) UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0 | Human actor should be warned to use only oref 0 OpenAPS should have an expiremental mode for human actor to play with other algorithms | The human is not informed that they should be using the latest oref0 | The human believes they understand the algorithm and that it's suitable for use | Human - Incomplete or inaccurate docs surrounding an algorithm
[6 empty] | Human - Human does not want to change working algorithm even though latest has more bug fixes | The OpenAPS should inform the human when there is a new update available |  |  | Human - Human actor doesn't get a notification email (UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0) | Whenever there is a improvement to the algorithm, human actor gets a notification email
[6 empty] | Human - Human is prototyping a new algorithm | OpenAPS system should come / start with test mode and the user should be aware that this is not a toy and documentation should outline major "Don'ts" as much as possible
[5 empty] | The human actor thinks the algorithm is suitable for use | Human - Human is not checking for algorithm updates | There should be e-mail notifications of available algorithm updates
| SO: Stop OpenAPS loop | UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices (eg when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations), but the OpenAPS loop is still issuing insulin doses | There is a problem with OpenAPS or OpenAPS is not able to make the insulin pump dose basal in line with its recommendations | Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS or when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations, though is informed OpenAPS cannot make the insulin pump give good insulin doses | The human actor believes there is a significant problem with the OpenAPS rig but continues to use it anyway | Human error (accident) UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices (eg when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations), but the OpenAPS loop is still issuing insulin doses | OpenAPS must inform the human actor whether it's recommendations are being applied OpenAPS must alarm the human actor if it suddenly stops or starts working | There is a significant problem with the OpenAPS rig, but the human is not informed of this | The human believes that the OpenAPS rig is working well | Human - Lack of feedback from OpenAPS (Human actor does not stop the loop when there is a problem with OpenAPS but OpenAPS is still issuing insulin doses) | OpenAPS should get feedback from the insulin pump when it has administered a dose, and should send this to the user Human actor should track blood glucose levels to make sure the OpenAPS treatment is working.
[6 empty] | Human - Human actor tries to stop the loop but it is not stopping | There should be feedback to the user indicating whether OpenAPS is successfully issuing dosing recommendations to the user |  |  | Human - OpenAPS rig does not 'realise' it's not working well | OpenAPS should get feedback from the insulin pump when it has administered a dose, and should send this to the user Human actor should track blood glucose levels to make sure the OpenAPS treatment is working.
[5 empty] | The human actor believes there is a problem with OpenAPS issuing doses to the insulin pump, but continues to use OpenAPS anyway |  |  | There is a problem with OpenAPS issuing doses to the insulin pump, but the human is not informed of this | The human believes that OpenAPS is able to issue doses to the insulin pump well | Human - OpenAPS does not pull feedback from insulin pump (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices, but the OpenAPS loop is still issuing insulin doses) | OpenAPS information should be verified and not compromised
[6 empty] | Human - Human keeps OpenAPS operational in case the doses start going through (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices) | Human should be able to override the guidance of OpenAPS and keep OpenAPS running |  |  | Human - OpenAPS rig does not 'realise' the insulin pump is not dosing as instructed | OpenAPS should get feedback from the insulin pump when it has administered a dose, and should send this to the user Human actor should track blood glucose levels to make sure the OpenAPS treatment is working.
The human actor does not believe there is a problem OpenAPS instructions to the insulin pump
|  | UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running | The OpenAPS system is expected to be up and running | Human actor provides Switch off (hold button goes into reset) when the system should be up and running, though is informed not to do this | The human actor believes the system is expected to be up and running, but provides switch off anyway | Human - Human actor suddenly needs to remove the insulin pump and OpenAPS rig | When the OpenAPS rig is disconnected, the insulin pump should fall back to normal operation | There is no reason for the system not to be up and running, but the human actor is not informed of this | The human actor believes that it is best to turn the system off | Human - OpenAPS algorithm has a bug which makes it indicate that the OpenAPS system needs switching off when it doesn't | OpenAPS system should supply a reason if it suggests that it needs to be turned off
[6 empty] | Human - Human actor wants to change something with the OpenAPS rig / insulin pump and switches off the OpenAPS rig | When human actor switches off the OpenAPS rig suddenly, the insulin pump goes back to normal operation and OpenAPS settings are retained as before |  |  | Human - Lack of feedback from rig or pump indicating that the OpenAPS looping is working (especially when dosing hasn't begun quite yet)
Human - Human actor knows of a hazard which OpenAPS doesn't
[6 empty] | Human - Human actor accidently turns off or disconnects the OpenAPS rig (battery falls out, wires caught, human takes rig out of range of pump etc) | If the OpenAPS rig is disconnected, the insulin pump should go back to regular operation
[5 empty] | The human actor does not believe the system is expected to be up and running | Human - Human actor mistinterprets a signal from the OpenAPS rig or pump or logs | There should be a forum and a resource which the human can use to understand the OpenAPS logs
[6 empty] | Human - Human actor believes the insulin dosing is wrong or not working (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) | Human actor needs to know what to do and how to check if the insulin dosing is working
[6 empty] | Human - OpenAPS is taking a while to begin and the human thinks this means it's not working properly | Include prompt or message in the OpenAPS UI that OpenAPS is still booting up; nothing is wrong.
[6 empty] | Human - Human confusion (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) | All of the buttons on the OpenAPS rig should be clearly labelled There should be feedback to the user indicating whether OpenAPS is successfully issuing dosing recommendations to the user
|  | UCA-SO-1000: Human actor breaks the Power Off button (or another part of the OpenAPS rig) | - | Human actor breaks the Power Off button, though is warned not to do this | The human actor believes that the power off button is vulnerable to breaking, but breaks it anyway | Human error (accident) (UCA-SO-1000: Human actor breaks the Power Off button (or another part of the OpenAPS rig)) | Any parts of the OpenAPS rig that should be treated with care should come with a warning | The Human actor is not warned of breaking the power button (or other OpenAPS rig part) | The human actor believes that the power button is not vulnerable to being broken | Human - There is no warning on the OpenAPS rig that OpenAPS is vulnerable to being broken | Any parts of an OpenAPS rig that should be treated with care should come with a warning
[5 empty] | The human actor does not believe that the power off button is vulnerable to breaking | Human - Human does not see warning | Warning should be easy to see for user and there should be a warning symbol permanently present when there is a problem |  |  | Human - Human does not see warning | Warning should be easy to see for user and there should be a warning symbol permanently present when there is a problem
| UD: Provide useful data to pump (eg carbs) | UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | The human actor has information which can improve insulin dosing recommendations | Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime), though is warned to do this | Human actor knows they have have information which can improve insulin dosing recommendations, but does not input this information | Human - Human is too busy or cannot be bothered to input data to the insulin pump | The human should have the option to avoid entering meal data | The human actor is not informed that it has data which can improve insulin dosing recomendations | The human actor does not believe she has data which can improve dosing recommendations | Human - Error in insulin pump instructions (UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | There should be a way for errors in the OpenAPS instructions to be reported Any safety critical input should trigger a prompt for verification
[6 empty] | Human error (accident) UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | OpenAPS should warn the human actor if human has not input useful daily information in X hours |  |  | Human - Incomplete insulin pump instructions | Any OpenAPS relevant gap in insulin pump documentation that is found by a user should be recorded in OpenAPS documentation or the OpenAPS forum
[5 empty] | Human actor does not realise they can input information to improve insulin dosing recommendations | Human - Human has not read the instructions of the insulin pump properly / read wrong instructions | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place. OpenAPS should give notification either when human entry to insulin pump looks incorrect or when human actor may have forgotten to enter info into the insulin pump. Human should still have ability to make decision (repeated requirement) |  |  | Human - Insulin pump doesn't accomodate for the useful data
??? | [5 empty] | Human - Human has read insulin pump instructions but does not know how to input data into current insulin pump in use | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place
[6 empty] | Human - Human forgets to input data or that they can input data which is useful | If mealtime data hasn't been entered into the insulin pump or OpenAPS rig for a long time, OpenAPS should estimate the meal data or prompt the human to enter their meal data
|  | UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing | (The data that the Human actor is providing to insulin pump is erroneous) | Human actor provides erroneous input of data affecting insulin dosing, though is informed not to do this | Human actor knows they should not input this data into the insulin pump, but does so anyway | Human - Human actor is experimenting |  | The human actor's information to inform a dosing recommendation is not accurate | The human actor has entered accurate information into the insulin pump | Human - Error in insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing) | There should be a way for errors in the OpenAPS instructions to be reported Any safety critical input should trigger a prompt for verification
[6 empty] | Human error (accident) UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing |  |  |  | Human - Incomplete insulin pump or OpenAPS instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely Any OpenAPS relevant gap in insulin pump documentation that is found by a user should be recorded in OpenAPS documentation or the OpenAPS forum
[5 empty] | Human actor believes they are not making an error in entering the data into the insulin pump | Human - Human actor has not read the insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing) | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place. OpenAPS should give notification either when human entry to insulin pump looks incorrect or when human actor may have forgotten to enter info into the insulin pump. Human should still have ability to make decision. |  |  | Human - Human actor is a different person to the user and doesn't know the accurate data of the user (parent doesn't know kid had some extra candy) | If the user data given to OpenAPS is causing recommendations which are leading to a non-optimal glucose level, OpenAPS should adjust it's dosing recommendation and remind the human user to check the calibration of the CGM
[6 empty] | Human - Human actor has not read or understood OpenAPS instructions for inputting data into the insulin pump | OpenAPS should not provide dosing recommendations if the insulin pump is not configured correctly for OpenAPS use |  |  | Human - Error in OpenAPS instructions | There should be a way for errors in the OpenAPS instructions to be reported
[6 empty] | Human - Human actor has read old / deprecated instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[6 empty] | Human - Instructions are complex and Human actor has forgetten them | OpenAPS instructions and any complexities should be as clear as possible
|  | UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime | Carb data is needed within X minutes of a meal | Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime, though is informed not to do this | The human actor knows they should should provide carb data to the pump within X minutes of a meal, but provides the data outside this time window anyway | Human error (accident) UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime | OpenAPS documentation should state a time limit in which to provide meal data | Carb data is needed within X minutes of a meal, but the human actor is not informed of this | The human actor does not believe Carb data is needed within X minutes of a meal | Human - Lack of information in docs (Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime) | Insulin pump or OpenAPS docs should recommend a timeframe to provide meal data to the insulin pump
[6 empty] | Human - Human trying to influence insulin dosing / glucose levels (UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime) | Make sure human knows the time for the insulin to react in their body and sets the advanced time (by X minutes) accordingly
[5 empty] | The human actor is not aware that they should provide OpenAPS meal data within X minutes of a meal | Human - Human does not read docs | OpenAPS safe defaults should be verified as part of each release
Set fallback basal rate on pump
| SB: Set bolus on pump | UCA-SB-100: Human actor sets a bolus (or basal) which unintentionally overrides OpenAPS temporary basals | The active OpenAPS algorithm is providing insulin dosing which should not be overriden | Human actor sets a bolus (or basal) which unintentionally overrides OpenAPS temporary basals, though is informed not to do this | The human actor knows the bolus will override OpenAPS temporary basal recommendations but demands it anyway | Human - Human user wants their bolus dose rather than the basal recommendation (if they have additional information about their current state or upcoming activities) |  | The human actor's bolus overrides the OpenAPS insulin dosing recommendations, but the human actor is not informed of this | The human actor does not believe her bolus request would override the OpenAPS dosing recommendation | Human - Error in OpenAPS instructions | There should be a way for errors in the OpenAPS instructions to be reported
[6 empty] | Human error (accident) UCA-SB-100: Human actor sets a bolus (or basal) which unintentionally overrides OpenAPS temporary basals |  |  |  | Human - Incomplete OpenAPS instructions | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
[5 empty] | The human actor does not know the bolus will override OpenAPS temporary basal recommendation | Human - Human actor has not read or understood the OpenAPS instructions | OpenAPS instructions must be clear and non-ambiguous |  |  | Human - Another old set of OpenAPS instructions is available
[10 empty] | Human - Human actor cannot find information | The OpenAPS docs must be easy to navigate, with a search bar
| PM: Provide settings mandatory for OpenAPS on pump | UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS | The insulin pump should be set as follows: - Temp Basal type as "units per hour" not "%" - Carb ratios to grams, not to exchange units - Max basal rate to reasonable value, normally no more than 4x regular basal - Basal profile, carb ratios and ISF values should be set (correctly) - DIA set high enough (should be 6 or 7 not 2 or 3 as most people use, OpenAPS defaults to 5 if this is not set) - The loop will not work if there are periods with basal settings of zero, instead set the lowest basal setting permitted and OpenAPS will issue temp basals of zero as needed | Human actor configures insulin pump settings in a way that is not suitable for OpenAPS, though is informed not to do this | The human actor believes the pump settings are not suitable for OpenAPS but uses them anyway | Human - Human actor is experimenting | OpenAPS system should come / start with test mode User should be aware that this is not a toy and documentation should outline major "Don'ts" as much as possible | The human actor's insulin pump settings are not suitable for OpenAPS, but the human actor is not informed of this | The human actor believes that they have configured the insulin pump suiltably for OpenAPS | Human - OpenAPS instructions have an error (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | There should be multiple reviews including signoffs and phase gates prior to release of OpenAPS instructions.
[6 empty] | Human error (accident) UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS | OpenAPS should give a warning when human actor has provided non recommended configurations |  |  | Human - OpenAPS instructions are incomplete (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | There should be multiple reviews including signoffs and phase gates prior to release of OpenAPS instructions.
[5 empty] | The human actor believes pump settings are suitable for OpenAPS | Human - Human actor believes the OpenAPS instructions are wrong | The human should be able to ask questions to an OpenAPS forum and preferably communicate with maintainers of the repository |  |  | Human - OpenAPS does not make sure that the insulin pump settings are suitable for OpenAPS to operate | OpenAPS should request the insulin pump settings / configurations and compare them to recommended values
Human - Human actor has read old / deprecated instructions
[6 empty] | Human - Human actor has not read or understood the OpenAPS instructions | OpenAPS instructions must be clear and non-ambiguous
[6 empty] | Human - Human actor has forgetten the OpenAPS instructions (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | OpenAPS should make a notification when the insulin pump settings do not match the recommended configuration in the OpenAPS documentation with X frequency
| UC: Update CGM / CGM receiver | UCA-UC-1: Human does not update CGM when there are updates available for X weeks which can improve glucose data accuracy or prevent problems
UCA-UC-100: Human updates the CGM when OpenAPS rig is trying to pull data from the CGM
UCA-UC-101: Human provides wrong CGM update

## Level 1 Requirements
| This column is normally hidden | For the moment, try not to change alphabetical order of the Loss Scenarios A limitation of this layout of the requirements sheet is we don't see the UCA next to the Loss Scenario and so can lack context of the meaning of the LS. To mitigate this, go to Loss Scenario page and paste the UCA in brackets after the Loss Scenario, that will be reflected here |  |  | If your Requirement is already a Controller Constraint, you can just write "Covered by Controller Constraint" in this column
|  |  | Loss Scenario | Loss Scenario Occurences | OpenAPS Requirement (Scenario Constraint) |  |  | Controller Constraints list (Also requirements)
OpenAPS rig does not have adequate data to make a dosing recommendation
Shaun | Flawed OpenAPS decision making algorithm (UCA-STB-1: OpenAPS does not set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation) |  | CGM clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | 2 | CGM clock must be accurate and match RasPi time and real clock time, to nearest X seconds All times should be alligned with NTP or a common time source | CC-STB-1 : OpenAPS must set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation |  | CC-CO-1: Human actor must update the configurations of OpenAPS accordingly if there is a change in the OpenAPS loop setup which affects the configurations
Shaun | OpenAPS rig hardware has not been set up correctly |  | CGM sends data to OpenAPS with the current timestamp rather than the time that the data was from | 1 | Timestamping needs to match with CGM and RasPi clock and real clock with X seconds Timestamp needs to be accurate time of the reading rather than the time of the data transfer | CC-STB-2: OpenAPS must set the temporary basal dose when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation |  | CC-CO-100: Human actor must not select an inaccurate option when configuring OpenAPS
Shaun | OpenAPS rig has run out of power |  | Data is sent from the insulin pump which is actually null data or dummy data | 2 | OpenAPS must validate data received from insulin pump |  |  | CC-CO-101: Human actor must not configure OpenAPS with a bad algorithm instead of oref0
Shaun | OpenAPS rig dependencies / software has been updated and OpenAPS now lacks compatibiliity with the insulin pump |  | Environment has noisy radio frequency or electromagnetic field | 1 | Communication between components must have an integrity check |  |  | CC-PM-100: Human actor must not configure insulin pump settings in a way that is not suitable for OpenAPS (Temp Basal rates as units per hour not %, carb ratios to grams not exchange units, max basal rate set too high, basal profile / carb ratios / ISF values not set, DIA too low)
Shaun | OpenAPS insulin dose recommendation algorithm is taking too long |  | Equipment failiure | 1 | There should be a heartbeat mechanism to monitor activeness between systems |  |  | CC-PTR-100: Human actor must not set up the OpenAPS rig wrong or insecurely (this needs more research, hardware will help)
Shaun | Flawed OpenAPS dosing calculation algorithm |  | Explorer board not sitting properly on the Edison | 1 | OpenAPS has a hardware integrity check and can communicate with all necessary peripherals | #REF! |  | CC-PTR-101: Human actor must not make a software update to another device (eg insulin pump, CGM) which makes the device incompatible with OpenAPS
Shaun | OpenAPS processes glucose data from Nightscout incorrectly |  | Extreme environmental changes | 1 | OpenAPS has a hardware integrity check and can communicate with all necessary peripherals There should be a heartbeat mechanism to monitor activeness between systems |  |  | CC-RG-1: OpenAPS must get valid data from the NightScout when it has not received data which is from the last X minutes
Shaun | OpenAPS processes data from the insulin pump incorrectly |  | Flawed OpenAPS data analysis / projection of future glucose level | 3 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-RG-100: OpenAPS must not request (and use) old or missing glucose data in place of recent glucose data
Shaun | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Flawed OpenAPS decision making algorithm (eg after pausing for a breakdown it does not resume issueing doses) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) | 1 | Covered by Controller Constraint |  |  | CC-RG-1000: OpenAPS must get data from NightScout at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
Shaun | OpenAPS rig is hacked by an imposter |  | Flawed OpenAPS decision making algorithm (eg is not designed to stop making recommendations / readjust its algorithms when the insulin pump is not in a state to give a dose) (UCA-STB-104 : OpenAPS makes a temporary basal recommendation when the insulin pump is not in a state to give a dose) | 1 | Covered by Controller Constraint |  |  | CC-RG-101: OpenAPS must not get data from the CGM more than X times a minute, if NightScout is not available (consequence: could diminish battery or wear out the CGM)
Shaun |  |  | Flawed OpenAPS decision making algorithm (either a bug or because the human engineer thinks that less information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-102: OpenAPS makes a temporary basal recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-RG-102: OpenAPS must request sufficient data from NightScout to make a reliable dosing recommendation
Shaun |  |  | Flawed OpenAPS decision making algorithm (either accidently or because the human engineer thinks that more information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-ROD-1: OpenAPS must get data from the insulin pump when it has not received data which is from the last X minutes
Shaun |  |  | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-RG-102: OpenAPS requests insufficient data from the Nightscout to make a reliable dosing recommendation) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-ROD-100: OpenAPS must not request (and use) old or missing data from the insulin pump in place of recent data (battery fail option)
Shaun |  |  | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-ROD-1000: OpenAPS must get data from the insulin pump at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
Shaun |  |  | Flawed OpenAPS decision making algorithm (UCA-RG-1: OpenAPS does not get data from Nightscout when it has not received data from the last X minutes) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-ROD-101: OpenAPS must not request data from the insulin pump more than X times a minute
Shaun | Flawed OpenAPS decision making algorithm (eg after pausing for a breakdown it does not resume issueing doses) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) |  | Flawed OpenAPS decision making algorithm (UCA-RG-100: OpenAPS requests missing glucose data) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely | CC-ROD-1: OpenAPS must get data from the insulin pump when it has not received data which is from the last X minutes |  | CC-ROD-102: OpenAPS must request sufficient data from the insulin pump to make a reliable dosing recommendation
Shaun | OpenAPS rig components are not connected together properly |  | Flawed OpenAPS decision making algorithm (UCA-RG-101: OpenAPS gets data from the CGM too frequently, if Nightscout is not available (consequence: could diminish battery of or wear out the CGM)) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely Controller constraint |  |  | CC-S-1: Human actor must install all the system dependencies necessary for OpenAPS
Shaun | OpenAPS rig has run out of power |  | Flawed OpenAPS decision making algorithm (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-S-100: Human actor must not make a software update to the software that the OpenAPS rig is using which makes it incompatible with other device(s) in the system (eg pump, CGM)
Shaun | OpenAPS rig dependencies / software has been updated and OpenAPS now lacks compatibiliity with the insulin pump |  | Flawed OpenAPS decision making algorithm (UCA-ROD-100: OpenAPS requests data from the insulin pump that's missing (battery fail option)) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely Controller constraint | CC-RG-1: OpenAPS must get valid data from the NightScout when it has not received data which is from the last X minutes |  | CC-S-101: Human actor must not install unecessary additional system dependencies for OpenAPS without a "test" configuration
Shaun | OpenAPS insulin dose calculation algorithm taking too long |  | Flawed OpenAPS decision making algorithm (UCA-ROD-101: OpenAPS requests data from the insulin pump too frequently, so that the pump is overloaded with requests (?)) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely Controller constraint |  |  | CC-S-2: Human actor must update system dependencies when there are compatible updates available
Shaun | OpenAPS has not been configured correctly |  | Flawed OpenAPS decision making algorithm (UCA-STB-1: OpenAPS does not set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely Controller constraint |  |  | CC-SB-100: Human actor must not set a bolus (or basal) which unintentionally permanently overrides OpenAPS temporary basals
Shaun | OpenAPS rig is hacked by an imposter |  | Flawed OpenAPS decision making algorithm (UCA-STB-103: OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) | 1 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely Config files should be verified when possible | #REF! |  | CC-SO-1: Human actor must stop the OpenAPS loop when there is a problem with OpenAPS or when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations
Shaun | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Flawed OpenAPS dosing calculation algorithm | 3 | OpenAPS algorithm should be tested previously and validated in realtime for accuracy and designed securely |  |  | CC-SO-100: Human actor must not provide Switch off (hold button goes into reset) when the system is expected to be up and running
Milan | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes its recommendation is still active when informed that it's not) |  | Human - Another old set of OpenAPS instructions is available | 1 | There should only be one set of official OpenAPS instructions, which contains all the information that needs to be known to use OpenAPS safely |  |  | CC-STB-1 : OpenAPS must set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation
Milan | Flawed OpenAPS decision making algorithm (either accidently or because the human engineer thinks that more information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-2: OpenAPS does not set the temporary basal rate when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation) |  | Human - Bug in dependencies script | 1 | The dependencies script must install all of the dependencies correctly | CC-S-1: Human actor must install all the system dependencies necessary for OpenAPS |  | CC-STB-100 : OpenAPS algorithm must not recommend a too high temporary basal dosing recommendation when current glucose levels are low and not likely to rise (based on active insulin and human insulin pump input)
Milan |  |  | Human - Docs out of date (Dependencies script out of date) | 1 | Whenever a change is made to the OpenAPS system or the compatible versions change, the docs should be updated | CC-S-2: Human actor must update system dependencies when there are compatible updates available |  | CC-STB-101 : OpenAPS algorithm must not recommend too low temporary basal dosing recommendation when current glucose levels are high and not likely to fall (based on current glucose levels, active insulin and human insulin pump input)
Milan |  |  | Human - Docs out of date (UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely) | 1 | Whenever a change is made to the OpenAPS system or the compatible versions change, the docs should be updated |  |  | CC-STB-102: OpenAPS must not make a temporary basal recommendation when there is inadequate and incomplete data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
Milan |  |  | Human - Docs out of date (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | 1 | Any errors that are found in the insulin pump instructions should be listed in the OpenAPS documentation There should be a way for errors in the insulin pump instructions to be directed to OpenAPS | CC-CO-1: Human actor must update the configurations of OpenAPS accordingly if there is a change in the OpenAPS loop setup which affects the configurations |  | CC-STB-103 : OpenAPS must not provide a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)
Milan | Flawed OpenAPS dosing calculation algorithm |  | Human - Docs out of date (when software has been updated) | 1 | Any errors that are found in the insulin pump instructions should be listed in the OpenAPS documentation There should be a way for errors in the insulin pump instructions to be directed to OpenAPS |  |  | CC-STB-104 : OpenAPS must not make a temporary basal recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
Milan | OpenAPS is using the wrong glucose data from Nightscout |  | Human - Docs out of date (working versions information) | 1 | There should be a way for errors in the OpenAPS instructions to be reported When OpenAPS software or working versions have changed, there should be a prompt or issue created to review and update the docs | CC-SO-1: Human actor must stop the OpenAPS loop when there is a problem with OpenAPS or when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations |  | CC-STB-2: OpenAPS must set the temporary basal dose when there is no active temporary basal recommendation and there is adequate data available to make a justified recommendation
Milan | OpenAPS does not account for human bolus dosing |  | Human - Error in insulin pump instructions (UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | 1 | There should be a way for errors in the OpenAPS instructions to be reported | #REF! |  | CC-UD-1: Human actor must provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime)
Milan | Flawed OpenAPS data analysis / projection of future glucose level |  | Human - Error in insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing) | 1 | There should be a way for errors in the OpenAPS instructions to be reported Any safety critical input should trigger a prompt for verification |  |  | CC-UD-100: Human actor must not provide erroneous input of data affecting insulin dosing
Milan | OpenAPS rig is hacked by an imposter |  | Human - Error in OpenAPS instructions | 5 | There should be a way for errors in the OpenAPS instructions to be reported | CC-UD-1: Human actor must provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) |  | CC-UD-1000: Human actor must provide meal data to the insulin pump within X minutes from the actual mealtime
Milan |  |  | Human - Error in OpenAPS instructions (Wrong version number of the software) | 1 | There should be a way for errors in the OpenAPS instructions to be reported |  |  | #REF!
Milan |  |  | Human - Flaw in the OpenAPS rig hardware setup design which the creators don't realise | 1 | OpenAPS rig hardware setup should be tested for robustness and fragiliity |  |  | #REF!
Milan |  |  | Human - Human actor accidently turns off or disconnects the OpenAPS rig (battery falls out, wires caught, human takes rig out of range of pump etc) | 1 | If the OpenAPS rig is disconnected, the insulin pump should go back to regular operation |  |  | #REF!
Milan |  |  | Human - Human actor believes the insulin dosing is wrong or not working (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) | 1 | Human actor needs to know what to do and how to check if the insulin dosing is working |  |  | #REF!
Milan | Flawed OpenAPS dosing calculation algorithm |  | Human - Human actor believes the OpenAPS instructions are wrong | 1 | The human should be able to ask questions to an OpenAPS forum and preferably communicate with maintainers of the repository |  |  | #REF!
Milan | OpenAPS is using the wrong glucose data from Nightscout |  | Human - Human actor cannot find information | 3 | The OpenAPS docs must be easy to navigate, with a search bar |  |  | #REF!
Milan | Flawed OpenAPS data analysis / projection of future glucose level |  | Human - Human actor cannot find information (installing dependencies) | 1 | The OpenAPS docs must be easy to navigate, with a search bar |  |  | #REF!
Milan | OpenAPS rig is hacked by an imposter |  | Human - Human actor cannot find information (UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly) | 1 | The OpenAPS docs must be easy to navigate, with a search bar |  |  | #REF!
Milan |  |  | Human - Human actor cannot find information (UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS) | 1 | The OpenAPS docs must be easy to navigate, with a search bar |  |  | #REF!
Milan |  |  | Human - Human actor cannot find information (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | 1 | The OpenAPS docs must be easy to navigate, with a search bar |  |  | #REF!
Milan |  |  | Human - Human actor cannot find information (working version) | 1 | OpenAPS should not provide dosing recommendations if it has not been tested to be safely compatible with the software versions of other components in the system |  |  | #REF!
Milan |  |  | Human - Human actor does not read the instructions about which software versions work | 1 | Human actor must get a notification e-mail when sofware versions change |  |  | #REF!
Milan |  |  | Human - Human actor doesn't get a notification email (UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0) | 1 | Whenever there is a improvement to the algorithm, human actor gets a notification email
Milan | OpenAPS is using old / missing / corrupt data |  | Human - Human actor has forgetten the OpenAPS instructions (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | 1 | OpenAPS should make a notification when the insulin pump settings do not match the recommended configuration in the OpenAPS documentation with X frequency
Milan | Flawed OpenAPS data analysis / projection of future glucose level |  | Human - Human actor has not read or understood OpenAPS instructions for inputting data into the insulin pump | 1 | OpenAPS should not provide dosing recommendations if the insulin pump is not configured correctly for OpenAPS use
Milan | Flawed OpenAPS decision making algorithm (either a bug or because the human engineer thinks that less information is needed to make a dosing recommendation than is actually necessary) (UCA-STB-102: OpenAPS makes a temporary basal recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)) |  | Human - Human actor has not read or understood the OpenAPS instructions | 2 | OpenAPS instructions must be clear and non-ambiguous
Milan |  |  | Human - Human actor has not read the insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing) | 1 | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place. OpenAPS should give notification either when human entry to insulin pump looks incorrect or when human actor may have forgotten to enter info into the insulin pump. Human should still have ability to make decision | CC-STB-100 : OpenAPS algorithm must not recommend a too high temporary basal dosing recommendation when current glucose levels are low and not likely to rise (based on active insulin and human insulin pump input)
Milan |  |  | Human - Human actor has read old / deprecated instructions | 2 | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely | CC-STB-101 : OpenAPS algorithm must not recommend too low temporary basal dosing recommendation when current glucose levels are high and not likely to fall (based on current glucose levels, active insulin and human insulin pump input)
Milan |  |  | Human - Human actor is a different person to the user and doesn't know the accurate data of the user (parent doesn't know kid had some extra candy) | 1 | If the user data given to OpenAPS is causing recommendations which are leading to a non-optimal glucose level... should OpenAPS adjust it's dosing recommendation or should it assume the glucose reading is incorrect? (Both to some extent) | CC-STB-102: OpenAPS must not make a temporary basal recommendation when there is inadequate and incomplete data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
Milan | Flawed OpenAPS decision making algorithm (UCA-STB-103: OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) |  | Human - Human actor is experimenting | 4 | OpenAPS system should come / start with test mode User should be aware that this is not a toy and documentation should outline major "Don'ts" as much as posisble | CC-STB-103 : OpenAPS must not provide a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)
Milan | Insulin pump misconfiguration is not fatal for OpenAPS, and OpenAPS is able to adjust its dosing to cause the correct glucose levels (UCA-STB-103 : OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) |  | Human - Human actor is rushing or thinks that they already know how to set up the rig | 1 | OpenAPS should send a warning to human actor when human actor hasn't set up the rig matching with the configurations | CC-STB-104 : OpenAPS must not make a temporary basal recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
Milan |  |  | Human - Human actor knows of a hazard which OpenAPS doesn't | 1 | Not a loss scenario | #REF!
Milan | Human has not entered some of the optional insulin pump settings and OpenAPS default values for insulin pump settings in preferences.json are not suitable for the user |  | Human - Human actor mistinterprets a signal from the OpenAPS rig or pump or logs | 1 | There should be a forum and a resource which the human can use to understand the OpenAPS logs
Milan | OpenAPS does not verify insulin pump configurations |  | Human - Human actor reads old / deprecated instructions | 4 | (There should only be one set of official instructions)
Milan |  |  | Human - Human actor suddenly needs to remove the insulin pump and OpenAPS rig | 1 | When the OpenAPS rig is disconnected, the insulin pump should fall back to normal operation
Milan |  |  | Human - Human actor tries to stop the loop but it is not stopping | 1 | There should be feedback to the user indicating whether OpenAPS is successfully issuing dosing recommendations to the user
Milan | #REF! |  | Human - Human actor wants to change something with the OpenAPS rig / insulin pump and switches off the OpenAPS rig | 1 | When human actor switches off the OpenAPS rig suddenly, the insulin pump goes back to normal operation and OpenAPS settings are retained as before
Milan | #REF! |  | Human - Human confusion (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) | 1 | All of the buttons on the OpenAPS rig should be clearly labelled There should be feedback to the user indicating whether OpenAPS is successfully issuing dosing recommendations to the user
Milan | #REF! |  | Human - Human does not check their notifications / e-mail | 1 | Human must be able to select their preferred frequencies of notifications
Milan | #REF! |  | Human - Human does not read docs | 3 | OpenAPS safe defaults should be verified as part of each release
Milan | #REF! |  | Human - Human does not see the SD card information in the documentation and uses the wrong SD card | 1 | OpenAPS should notify user when SD card is of the wrong size or format Important information in docs should be easy to find in the hardware setup area of the docs
Milan | OpenAPS is making recommendations in case the insulin pump comes back online |  | Human - Human does not see warning | 2 | Warning should be easy to see for user and there should be a warning symbol permanently present when there is a problem | CC-ROD-100: OpenAPS must not request (and use) old or missing data from the insulin pump in place of recent data (battery fail option)
Milan | Flawed OpenAPS decision making algorithm (eg is not designed to stop making recommendations / readjust its algorithms when the insulin pump is not in a state to give a dose) (UCA-STB-104 : OpenAPS makes a temporary basal recommendation when the insulin pump is not in a state to give a dose) |  | Human - Human does not want to change working algorithm even though latest has more bug fixes | 1 | The OpenAPS should inform the human when there is a new update available | CC-ROD-101: OpenAPS must not request data from the insulin pump more than X times a minute
Milan | OpenAPS does not verify whether the insulin pump is online and delivering the insulin doses |  | Human - Human error (Human actor defaults to installing the latest software and dependencies and doesn't consider that this could cause a failiure) | 1 | OpenAPS docs should advise the human to only install dependencies via dependencies script (or at least see the versions here) and not manually | CC-ROD-102: OpenAPS must request sufficient data from the insulin pump to make a reliable dosing recommendation
Milan | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes that the insulin pump is operating and dosing properly when it is not |  | Human - Human forgets about the contents of the configurations or that they need to update the configs at all | 1 | When there is a change to the system setup, OpenAPS should detect this and prompt the human to update the configs | CC-RG-100: OpenAPS must not request (and use) old or missing glucose data in place of recent glucose data
Milan |  |  | Human - Human forgets to input data or that they can input data which is useful | 1 | If mealtime data hasn't been entered into the insulin pump or OpenAPS rig for a long time, OpenAPS should estimate the meal data or prompt the human to enter their meal data | CC-RG-101: OpenAPS must not get data from the CGM more than X times a minute, if NightScout is not available (consequence: could diminish battery or wear out the CGM)
Milan |  |  | Human - Human has not read the instructions of the insulin pump properly / read wrong instructions | 1 | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place. OpenAPS should give notification either when human entry to insulin pump looks incorrect or when human actor may have forgotten to enter info into the insulin pump. Human should still have ability to make decision (repeated requirement) | CC-RG-102: OpenAPS must request sufficient data from NightScout to make a reliable dosing recommendation
Milan |  |  | Human - Human has read insulin pump instructions but does not know how to input data into current insulin pump in use | 1 | OpenAPS documentation should make it clear that the human will be able to operate the pump manually, before the human puts the loop in place. | #REF!
Milan |  |  | Human - Human is not checking for algorithm updates | 1 | There should be e-mail notifications of available algorithm updates | CC-PTR-100: Human actor must not set up the OpenAPS rig wrong or insecurely (this needs more research, hardware will help)
Milan |  |  | Human - Human is not notified when there are updates to OpenAPS are available | 1 | There should be e-mail notifications of available OpenAPS updates | CC-PTR-101: Human actor must not make a software update to another device (eg insulin pump, CGM) which makes the device incompatible with OpenAPS
Milan |  |  | Human - Human is not prompted to update system dependencies after the initial install (?) | 1 | There should be an e-mail or OpenAPS prompt to update dependencies | CC-S-100: Human actor must not make a software update to the software that the OpenAPS rig is using which makes it incompatible with other device(s) in the system (eg pump, CGM)
Milan |  |  | Human - Human is prototyping a new algorithm | 1 | OpenAPS system should come / start with test mode Or user should be aware that this is not a toy and documentation should outline major "Don'ts" as much as possible | CC-S-101: Human actor must not install unecessary additional system dependencies for OpenAPS without a "test" configuration
Milan |  |  | Human - Human is too busy or cannot be bothered to input data to the insulin pump | 1 | The human should have the option to avoid entering meal data
Milan |  |  | Human - Human keeps OpenAPS operational in case the doses start going through (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices) | 1 | Human should be able to override the guidance of OpenAPS and keep OpenAPS running | CC-CO-100: Human actor must not select an inaccurate option when configuring OpenAPS
Milan | OpenAPS is using a flawed algorithm which is causing too frequent bolus recommendations |  | Human - Human misreads compatible version numbers | 1 | OpenAPS should check that human is using compatible versions of components. New versions of components should be tested | CC-CO-101: Human actor must not configure OpenAPS with a bad algorithm instead of oref0
Milan |  |  | Human - Human misunderstanding (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) | 1 | Covered by Controller Constraint (Human actor must not install unecessary additional dependencies for OpenAPS without a "test" configuration) | CC-SO-100: Human actor must not provide Switch off (hold button goes into reset) when the system is expected to be up and running
Milan | OpenAPS processing delays |  | Human - Human thinks that they already know how to set up the rig and don't check instructions (UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS) | 1 | Add prompts in OpenAPS UI that insists that human reviews and understands instructions prior to/during/after setting up the rig. Nice to have is an educational video on how to set up the rig | #REF!
Milan | Linux kernel scheduling has caused OpenAPS to dose bolus later than X minutes after the meal |  | Human - Human trying to influence insulin dosing / glucose levels (UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime) | 1 | Make sure human knows the time for the insulin to react in their body and sets the advanced time (by X minutes) accordingly (Note, Is it possible for device to be 'conversive' and react to changes in glucose levels with questions about recent human activity, also note fast acting insulin is preferred)
Milan |  |  | Human - Human user wants their bolus dose rather than the basal recommendation (if they have additional information about their current state or upcoming activities) | 1 | Not a loss scenario | CC-UD-100: Human actor must not provide erroneous input of data affecting insulin dosing
Milan |  |  | Human - Incomplete docs | 1 | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely | CC-SB-100: Human actor must not set a bolus (or basal) which unintentionally permanently overrides OpenAPS temporary basals
Milan |  |  | Human - Incomplete instructions | 3 | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely | CC-PM-100: Human actor must not configure insulin pump settings in a way that is not suitable for OpenAPS (Temp Basal rates as units per hour not %, carb ratios to grams not exchange units, max basal rate set too high, basal profile / carb ratios / ISF values not set, DIA too low)
Milan | Flawed OpenAPS decision making algorithm (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes) |  | Human - Incomplete insulin pump instructions | 1 | Any OpenAPS relevant gap in insulin pump documentation that is found by a user should be recorded in OpenAPS documentation or the OpenAPS forum
Milan | OpenAPS only pulls data from the insulin pump when it detects changes on the insulin pump (not a problem, just a different method) |  | Human - Incomplete insulin pump or OpenAPS instructions | 1 | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely Any OpenAPS relevant gap in insulin pump documentation that is found by a user should be recorded in OpenAPS documentation or the OpenAPS forum
Milan | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Human - Incomplete OpenAPS instructions | 1 | There should only be one set of official OpenAPS instructions online, which contains all the information that needs to be known to use OpenAPS safely
Milan | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled |  | Human - Incomplete or inaccurate docs surrounding an algorithm | 1 | (see above)
Milan |  |  | Human - Incomplete or inaccurate docs, missing out a configuration setting | 2 | (see above)
Milan |  |  | Human - Instructions are complex and Human actor has forgetten them | 1 | OpenAPS instructions and any complexities should be as clear as possible
Milan | Flawed OpenAPS decision making algorithm (UCA-ROD-100: OpenAPS requests data from the insulin pump that's missing (battery fail option)) |  | Human - Insulin pump doesn't accomodate for the useful data | 1 | (No requirement, this is beyond spec of OpenAPS to deal with all possible data, but OpenAPS should help to accomodate these extra factors)
Milan | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | RESUME ME | Human - Lack of feedback from OpenAPS (Human actor does not stop the loop when there is a problem with OpenAPS but OpenAPS is still issuing insulin doses) | 1 | OpenAPS should get feedback from the insulin pump when it has administered a dose, and should send this to the user Human actor should track blood glucose levels to make sure the OpenAPS treatment is working.
Milan | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Human - Lack of feedback from rig or pump indicating that the OpenAPS looping is working (especially when dosing hasn't begun quite yet) | 1 | OpenAPS should get feedback from the insulin pump when it has administered a dose, and should send this to the user Human actor should track blood glucose levels to make sure the OpenAPS treatment is working.
Milan |  |  | Human - Lack of information in docs (Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime) | 1 | Insulin pump or OpenAPS docs should recommend a timeframe to provide meal data to the insulin pump
Milan |  |  | Human - New insulin pump software has not been properly tested with OpenAPS | 1 | Add configuration check between systems including of their software/firmware versions used and only allow system to work with insulin pumps that have been properly tested
Milan |  |  | Human - New OpenAPS software has not been properly tested with all the other devices (insulin pumps, CGM) | 1 | (see above)
Milan | Flawed OpenAPS decision making algorithm (UCA-ROD-101: OpenAPS requests data from the insulin pump too frequently, so that the pump is overloaded with requests (?)) |  | Human - Nobody at all is checking whether there are updates to OpenAPS system dependencies available | 1 | The versions of OpenAPS dependencies that are installed should be fixed Somebody should keep a look out and test for major updates to OpenAPS dependencies
Milan |  |  | Human - Nobody monitors and tests the latest updates to determine if they are compatible with OpenAPS | 1 | Major updates to OpenAPS should be tested
Milan |  |  | Human - OpenAPS algorithm has a bug which makes it indicate that the OpenAPS system needs switching off when it doesn't | 1 | OpenAPS system should supply a reason if it suggests that it needs to be turned off
Milan |  |  | Human - OpenAPS does not make sure that the insulin pump settings are suitable for OpenAPS to operate | 1 | OpenAPS should request the insulin pump settings / configurations and compare them to recommended values
Milan |  |  | Human - OpenAPS does not pull feedback from insulin pump (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices, but the OpenAPS loop is still issuing insulin doses) | 1 | OpenAPS information should be verified and not compromised
Milan | OpenAPS believes that no other relevant information is available to be requested when there is |  | Human - OpenAPS gives the wrong information to the human (possible wrong version number, possible multiple instances of same program etc.) | 2 | OpenAPS docs regarding other components should be reviewed regularly, in keeping with the latest recommended approach
Milan | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the insulin pump (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) |  | Human - OpenAPS instructions are incomplete (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | 1 | There should be multiple reviews including signoffs and phase gates prior to release of OpenAPS instructions.
Milan | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) |  | Human - OpenAPS instructions have an error (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS) | 1 | There should be multiple reviews including signoffs and phase gates prior to release of OpenAPS instructions. | #REF!
Milan |  |  | Human - OpenAPS is taking a while to begin and the human thinks this means it's not working properly | 1 | Include prompt or message in the OpenAPS UI that OpenAPS is still booting up; nothing is wrong.
Milan |  |  | Human - OpenAPS rig does not 'realise' it's not working well | 1 | OpenAPS should get feedback from the insulin pump when it has administered a dose
Milan |  |  | Human - OpenAPS rig does not 'realise' the insulin pump is not dosing as instructed | 1 | OpenAPS should get feedback from the insulin pump when it has administered a dose
Milan |  |  | Human - There is no warning on the OpenAPS rig that OpenAPS is vulnerable to being broken | 1 | Any parts of an OpenAPS rig that should be treated with care should come with a warning
Milan | OpenAPS requests information at an independent time to giving dosing recommendations |  | Human - There's a new potential device that hasn't been included in the configs | 1 | Developers should try and involve any new device which would improve the functionality of OpenAPS
Milan | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Human - Vulnerability to malicious actor interacting | 2 | OpenAPS follows good secure coding practices OpenAPS rig should ask for verification when the human actor has requested information
Milan | The time that it takes for OpenAPS to make a dosing recommendation from receiving information has been lengthened, but the OpenAPS algorithm has not been adjusted for this |  | Human - Vulnerability with one of the dependencies / software that OpenAPS people are unaware of | 1 | OpenAPS should use only well trusted and well tested dependencies
Milan | Poor task scheduling from OpenAPS, - operating system, badly implemented call-flow |  | Human actor has input too few carbs or gives incorrect information relating to exercise / other factors | 1 | OpenAPS should give a warning when human actor has provided abnormal carb data
Milan | Poor task scheduling from OpenAPS - wrong order algorithm, the dosing recommendation is always calculated before reading data and the recommendation is always made on the base of old data |  | Human actor has input too many carbs or gives incorrect information relating to exercise / other factors | 1 | OpenAPS should give a warning when human actor has provided abnormal carb data
Milan |  |  | Human error (accident) (UCA-S-1: Human actor does not install all the system dependencies necessary for OpenAPS) | 1 | OpenAPS loop should not run without necessary dependencies being installed | CC-ROD-1000: OpenAPS must get data from the insulin pump at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
Milan | Flawed OpenAPS decision making algorithm (UCA-RG-1: OpenAPS does not get data from Nightscout when it has not received data from the last X minutes) |  | Human error (accident) (UCA-S-2: Human actor does not update OpenAPS system dependencies when there are compatible updates available) | 1 | When new updates have been tested to be compatible, OpenAPS should issue an email notification
Milan | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Human error (accident) (UCA-SO-1000: Human actor breaks the Power Off button (or another part of the OpenAPS rig)) | 1 | Any parts of the OpenAPS rig that should be treated with care should come with a warning
Milan | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled |  | Human error (accident) UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly | 1 | Where possible, OpenAPS should detect a change in the loop setup and request that the human actor updates the configurations | CC-RG-1000: OpenAPS must get data from NightScout at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
Milan |  |  | Human error (accident) UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS | 1 | When configuring OpenAPS, human actor should be asked to verify decisions
Milan |  |  | Human error (accident) UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0 | 1 | Human actor should be warned to use only oref 0 OpenAPS should have an expiremental mode for human actor to play with other algorithms
Milan |  |  | Human error (accident) UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS | 1 | OpenAPS should give a warning when human actor has provided non recommended configurations | #REF!
Milan | Flawed OpenAPS decision making algorithm (UCA-RG-100: OpenAPS requests missing glucose data) |  | Human error (accident) UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely | 1 | OpenAPS rig setup instructions should be clear, complete, easy to find and cover all possible hardware
Milan | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled |  | Human error (accident) UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS | 1 | When new updates have been tested to be compatible, OpenAPS should issue an email notification OpenAPS should list compatible versions in the documentation
Milan | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Human error (accident) UCA-S-100: Human actor makes a software update to the software that OpenAPS is using, which makes it incompatible with other device(s) in the system (eg pump, CGM) | 1 | When new updates have been tested to be compatible, OpenAPS should issue an email notification OpenAPS should list compatible versions in the documentation
Milan |  |  | Human error (accident) UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS | 1 | Human actor should be discouraged from installing any other dependencies that are not in the installation script
Milan |  |  | Human error (accident) UCA-SB-100: Human actor sets a bolus (or basal) which unintentionally overrides OpenAPS temporary basals | 1 | Don't think this is possible with oref0
Milan |  |  | Human error (accident) UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices (eg when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations), but the OpenAPS loop is still issuing insulin doses | 1 | OpenAPS must inform the human actor whether it's recommendations are being applied OpenAPS must alarm the human actor if it suddenly stops or starts working
Milan | Flawed OpenAPS decision making algorithm (UCA-RG-101: OpenAPS gets data from the CGM too frequently, if Nightscout is not available (consequence: could diminish battery of or wear out the CGM)) |  | Human error (accident) UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) | 1 | OpenAPS should warn the human actor if human has not input useful daily information in X hours
Milan |  |  | Human error (accident) UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing | 1 | When providing data input, human actor should be asked to verify decisions
Milan |  |  | Human error (accident) UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime | 1 | OpenAPS documentation should state a time limit in which to provide meal data | #REF!
Milan |  |  | Human has not entered some of the optional insulin pump settings and OpenAPS default values for insulin pump settings in preferences.json are not suitable for the user | 1 | Documentation must clearly request the human actor how to correctly test and apply all the insulin pump settings which would be necessary to give good dosing recommendations
Milan |  |  | Human suddenly consumes a lot of food / sugar | 1 | Human actor must not suddenly consume a lot of food / sugar OpenAPS must warn the human actor if insulin is in low supply | CC-UD-1000: Human actor must provide meal data to the insulin pump within X minutes from the actual mealtime
Shaun | OpenAPS believes that no other relevant information is available to be requested when there is |  | Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | 3 | OpenAPS should notify user when insulin pump clock time differs from RasPi system clock
Shaun | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the Nightscout |  | Insulin pump falls back to standard operation (not using OpenAPS) | 2 | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user.
Shaun | Flawed OpenAPS decision making algorithm (perhaps not taking into account the possibility of null data or conflicting data) (UCA-RG-102: OpenAPS requests insufficient data from the Nightscout to make a reliable dosing recommendation) |  | Insulin pump has received an operating data request from another party (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes ) | 1 | OpenAPS should only share information to authorised individuals |  | Note: There could be legitimate use cases for this (eg parent makes the change)
Shaun |  |  | Insulin pump has run out of insulin | 1 | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user. OpenAPS must alert human actor if the insulin pump has stopped giving feedback
Shaun |  |  | Insulin pump has run out of power | 2 | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user. OpenAPS must alert human actor if the insulin pump has stopped giving feedback
Shaun |  |  | Insulin pump is blocked physically from the OpenAPS rig by another object | 1 | Insulin pump and OpenAPS rig must not be blocked off physically to the extent that they cannot connect, and this should be explained in the docs
Shaun | OpenAPS requests information independently of giving dosing recommendations |  | Insulin pump is configured with the wrong DIA / ISF / carb ratio settings | 2 | OpenAPS docs must advise human actor how they can determine the correct settings for DIA / ISF / carb ratio OpenAPS should advise human actor if something seems incorrect with the insulin pump settings
Shaun | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) |  | Insulin pump is not aware that it is not in a state to make a dosing recommendation | 3 | OpenAPS should get feedback from insulin pump about whether the pump has applied the dosing recommendation from OpenAPS
Shaun | The time that it takes for OpenAPS to make a dosing recommendation from receiving information has been lengthened, but the OpenAPS algorithm has not been adjusted for this |  | Insulin pump is not operating correctly but OpenAPS is not informed of this | 1 | OpenAPS should get feedback from insulin pump about whether the pump has applied the dosing recommendation from OpenAPS
Shaun | Poor task scheduling from OpenAPS, - operating system, badly implemented call-flow |  | Insulin pump misconfiguration is not fatal for OpenAPS, and OpenAPS is able to adjust its dosing to cause the correct glucose levels (UCA-STB-103 : OpenAPS provides a temporary basal dosing rate when the insulin pump settings are not configured correctly for OpenAPS (*1)) | 1 | -
Shaun | Poor task scheduling from OpenAPS - wrong order algorithm, the dosing recommendation is always calculated before reading data and the recommendation is always made on the base of old data |  | Insulin pump scheduling causes commands from OpenAPS to be executed too late | 1 | -
Jason |  |  | Insulin pump sends data to OpenAPS with the current timestamp rather than the time that the data was from | 1 | OpenAPS shall have a time expectation of when it receives data from the insulin pump and compare the time expectation with the timestamp - plausibility check
Jason | Human error (accident) UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely |  | Insulin pump sends data without including any indicator that the data is not valid, recent data | 1 | OpenAPS plausibility check shall examine raw data received from the insulin pump to ensure that the data has not been corrupted
Jason | Human - Human actor is experimenting |  | Insulin pump uses different units | 1 | OpenAPS plausibility check shall include examination of the raw data from the insulin pump to ensure the data is in its expected range
Jason |  |  | Insulin software / firmware has been updated and insulin pump now lacks compatibiliity with OpenAPS | 1 | OpenAPS must verify that the insulin pump software and firmware is compatible with OpenAPS OpenAPS rig or docs should be able to indicate which insulin pump software versions have been tested
Jason | Human - Human actor is rushing or thinks that they already know how to set up the rig |  | Lack of algorithm programmer understanding of the information level required to make a dosing recommendation | 3 | Code inspection
Jason | Human - Human actor reads old / deprecated instructions |  | Lack of glucose data verification | 1 | Human actor must verify the CGM readings every X hours with a glucose meter.
Jason | Human - Human does not see the SD card information in the documentation and uses the wrong SD card |  | Linux kernel scheduling has caused OpenAPS to dose bolus later than X minutes after the meal | 1 | OpenAPS shall employ a monitoring task to ensure that all safety-related threads in the application are triggered by the kernel on a timely basis (e.g. watchdog timer)
Jason |  |  | Liquid in the tubing connector has caused the infusion set to get blocked | 1 | ? need a way to detect this, either from insulin pump warning or other measure, and warn the user ?
Jason |  |  | Nightscout / CGM software / firmware has been updated and it now lacks compatibiliity with OpenAPS | 1 | OpenAPS shall perform compatibility checks with all interfacing applications and software and warn the human actor if there is an incompatibility detected
Jason |  |  | Nightscout has received an operating data request from another party | 1 | OpenAPS shall receive warning from NightScout if there is an operating request from another party and warn the human actor
Jason | Human error (accident) UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS |  | Nightscout is not online | 1 | OpenAPS shall warn the human actor if NightScout is not online | #REF!
Jason | Human - Human misreads compatible version numbers |  | Nightscout sends data to OpenAPS with the current timestamp rather than the time that the data was from | 1 | OpenAPS shall have a time expectation of when it receives data from NightScout and compare the time expectation with the timestamp - plausibility check
Jason | Human - Human actor reads old / deprecated instructions |  | OpenAPS assumes that since data has just been pulled it is new data, when actually new data was not available and null, old or corrupt data was pulled | 4 | OpenAPS shall have time expectation, plausiblity check, and check of timestamp to ensure validity of data
Jason | Human - Human thinks that they already know how to set up the rig and don't check instructions (UCA-PTR-101: Human actor makes a software update to another device (eg insulin pump, CGM), which makes it incompatible with OpenAPS) |  | OpenAPS believes that no other relevant information is available to be requested when there is | 2 | OpenAPS must check every X minutes for all the latest information OpenAPS must not use information older than X minutes for dosing recommendation
Jason |  |  | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes its recommendation is still active when informed that it's not) | 1 | If OpenAPS has incomplete data, the insulin pump should use the human's programmed basal rates
Jason |  |  | OpenAPS bug causes it to process data from the insulin pump incorrectly such that OpenAPS believes that the insulin pump is operating and dosing properly when it is not | 1 | OpenAPS must be aware of the correct IOB and active dosing (including OpenAPS must be made aware if the insulin pump / has been disconnected from the body or OpenAPS)
Jason |  | NOTE: Change to System clock | OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time) | 9 | OpenAPS shall be timesynced with all required components and require the human actor to update the time if it is necessary ? is this possible?
Jason | Human error (accident) (UCA-S-1: Human actor does not install all the system dependencies necessary for OpenAPS) |  | OpenAPS does not account for human bolus dosing | 1 | ? this has to be human actor responsibility. OpenAPS could take action if it does see an unexpected change due to the bolus but at that point it may be too late ?
Jason | Human - Human actor reads old / deprecated instructions |  | OpenAPS does not know that the human has missed a meal / unusual circumstances / travelling | 1 | ? this also has to be human actor responsibility. See row above ?
Jason |  |  | OpenAPS does not pull all of the relevant insulin pump configuration information from the insulin pump | 1 | OpenAPS should prompt the human for data if glucose levels are X mg lower than expected
Jason |  |  | OpenAPS does not verify insulin pump configurations | 1 | OpenAPS shall have a monitoring task that checks whether or not OpenAPS has verified the insulin pump configuration and warn the human actor if this has not been done (and/or safe shutdown).
Jason |  |  | OpenAPS does not verify whether the insulin pump is online and delivering the insulin doses | 1 | OpenAPS shall have a monitoring task (watchdog timer) that ensures that the OpenAPS task that verifies that the insulin pump is online is executing correctly.
Jason |  |  | OpenAPS has gone out of range from Nightscout and not received updated glucose information | 1 | OpenAPS shall warn the human actor if it is no longer connected to and/or no longer receiving information from NightScout. There should be checks for wireless interference every X minutes, and OpenAPS should send a warning to the user if there is a positive result
Jason | Human error (accident) (UCA-S-2: Human actor does not update OpenAPS system dependencies when there are compatible updates available) |  | OpenAPS has gone out of range from the insulin pump and cannot issue a dosing recommendation OpenAPS has had wireless interference from the insulin pump and cannot issue a dosing recommendation | 1 | See row 163
Jason | Human - Human does not check their notifications / e-mail |  | OpenAPS has gone out of range from the insulin pump and not received updated operating information | 1 | See row 163
Jason |  |  | OpenAPS has not been configured correctly | 1 | OpenAPS shall have a monitoring task that continuously checks for correct and/or plausible system configuration. ? additional configuration checks may need to be implemented ?
Jason |  |  | OpenAPS has not received correct data from the human actor which would conflict with the rest of the current data and prevent a recommendation | 1 | OpenAPS shall warn the human actor if it detects data from the human actor that is implausible or in conflict with the current state of the system. ? if this continues, OpenAPS may need to consider safe shutdown ? | #REF!
Jason |  |  | OpenAPS has received inaccurate glucose data from Nightscout | 5 | OpenAPS must make human verify data
Jason | Human error (accident) UCA-S-100: Human actor makes a software update to the software that OpenAPS is using, which makes it incompatible with other device(s) in the system (eg pump, CGM) |  | OpenAPS has received inaccurate glucose data from the human to override glucose data from Nightscout | 2 | OpenAPS should prompt and possibly force glucose monitor information to be consistently tested against the human
Jason | Human - Human error (Human actor defaults to installing the latest software and dependencies and doesn't consider that this could cause a failiure) |  | OpenAPS has received incorrect data from the human actor which conflicts with the rest of the current data and prevents a recommendation | 1 | See row 167
Jason | Human - Human actor reads old / deprecated instructions |  | OpenAPS insulin dose calculation algorithm taking too long | 1 | OpenAPS application shall have a monitoring task that will check if tasks are taking too long to execute and warn the user if this occurs.
Jason | Human - Human actor does not read the instructions about which software versions work |  | OpenAPS insulin dose recommendation algorithm is taking too long | 1 | The time between OpenAPS being sent the data to make the dosing calculation and OpenAPS issuing and insulin dose must be < X seconds
Jason |  |  | OpenAPS is having problems communicating with insulin pump | 1 | Communication must be safe guarded Irregularities in the communication must be detected and actions/safe states used during communication problems
Jason |  |  | OpenAPS is making recommendations in case the insulin pump comes back online | 1 | Communication interruption must be detected Re-establishing of communication must be done in a defined way (make sure there is no stale data in the pipeline)
Jason | Human error (accident) UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS |  | OpenAPS is not informed of human bolus dosing | 1 | There must be a clear warning indication, if bolus dosing is not set and so not used in the calculation.
Jason | Human - Human actor is experimenting |  | OpenAPS is not informed of other processes in the schedule | 2 | Process scheduling in the OS must ensure enough time for OpenAPS to run its functions | #REF!
Jason | Human - Human misunderstanding (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS) |  | OpenAPS is using a flawed algorithm which is causing too frequent bolus recommendations | 1 | Correct timing of bolus recommendations must be defined and monitored.
Jason |  |  | OpenAPS is using old / missing / corrupt data | 1 | Validity of used data must be checked. Only valid data shall be used for the openAPS algorithms.
Nicole |  |  | OpenAPS is using the wrong glucose data from Nightscout | 2 | Validity of used glucose data from Nightscout must be checked. (Generally data coming from NightScout should maybe be checked?) Only valid data shall be used for the openAPS algorithms. Nightscout data shall be checked regarding: - plausibility of its data in comparison with the valid ranges approved by the FDA (e.g. the ones used be the pump and glucose monitor) - plausibility of state changes - data on Nightscout has not been corrupted and has integrity checks
Nicole | Human error (accident) UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly |  | OpenAPS may lack some information that conflicts with information that it has, which means it should hold back on making a decision | 1 | Data must be checked for plausibilty and validity
Nicole | Human - Human forgets about the contents of the configurations or that they need to update the configs at all |  | OpenAPS only pulls data from the insulin pump when it detects changes on the insulin pump (not a problem, just a different method) | 1 | -
Nicole | Human - Human does not read docs |  | OpenAPS processes data from the insulin pump incorrectly | 1 | Processing of critical data needs to be tested and monitored
Nicole |  |  | OpenAPS processes glucose data from Nightscout incorrectly | 1 | Processing of critical data needs to be tested and monitored
Nicole |  |  | OpenAPS processing delays | 1 | OS scheduler must ensure correct processing processing of critical data needs to be monitored
Nicole | Human error (accident) UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS |  | OpenAPS radio frequency not tuned correctly | 1 | communication needs to be monitored there shall be some feedback from OpenAPS if communication is flawed
Nicole | Human - Human does not read docs |  | OpenAPS radio stick is not plugged into the OpenAPS rig properly | 1 | User manual needs to specify that plugged in connections must be checked for correct connection Communication needs to be monitored (is there a possibility of feedback from the glucose monitor or the pump?)
Nicole |  |  | OpenAPS request is designed for a different style of insulin pump or different insulin pump software | 1 | Compatible setups must be specified OpenAPS should give a warninig if it is connected to incompatible devices or software versions in these devices
Nicole |  |  | OpenAPS request is designed for a different version of Nightscout | 1 | Compatible setups must be specified OpenAPS should give a warninig if it is connected to an incompatible version of NightScout
Nicole |  |  | OpenAPS request is formatted erroneously | 2 | Compatible setups must be specified Correct openAPS request formatting must be specified, verified and monitored
Nicole | Human error (accident) UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0 |  | OpenAPS requests information at an independent time to giving dosing recommendations | 1 | Correct openAPS request timing must be specified, verified and monitored
Nicole | Human - Human does not want to change working algorithm even though latest has more bug fixes |  | OpenAPS requests information independently of giving dosing recommendations | 1 | Correct openAPS request timing must be specified, verified and monitored
Nicole | Human - Human is prototyping a new algorithm |  | OpenAPS rig components are not connected together properly | 1 | User manual needs to specify that plugged in connections must be checked for correct connection Communication/power supplies needs to be monitored and trigger an alarm
Nicole | Human - Human is not checking for algorithm updates |  | OpenAPS rig dependencies / software has been updated and OpenAPS now lacks compatibiliity with the insulin pump | 2 | Compatible setups must be specified When backward compatibility is affected by an update, it needs to be clearly flagged Q: NEED TO CHECK IF THIS IS ALSO SPECIFIED FOR NIGHTSCOUT
Nicole |  |  | OpenAPS rig does not have adequate data to make a dosing recommendation | 1 | requesting/receiving of critical data needs to be monitored Insulin pump should fall back to regular non-OpenAPS operation
Nicole | Human error (accident) UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices (eg when OpenAPS is not able to make the insulin pump dose basal in line with its recommendations), but the OpenAPS loop is still issuing insulin doses |  | OpenAPS rig hardware has not been set up correctly | 1 | It must be specified that timestamps of all connected devices need to match and how to verifiy this
Nicole | Human - Human actor tries to stop the loop but it is not stopping |  | OpenAPS rig has differing timestamps to the insulin pump and believes the last recommendation was made recently | 2 | User manual must specify the correct setup Software must check all required dependencies are present Hardware connections need to be set up to match the configuration the human entered to the openAPS configuration.
Nicole |  |  | OpenAPS rig has run out of power | 2 | Power supply must be monitored and low power supply must trigger an alarm How to avoid low power supply must be defined in the user manual
Nicole | Human - Human keeps OpenAPS operational in case the doses start going through (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices) |  | OpenAPS rig is hacked by an imposter | 4 | The OpenAPS rig and it's potentially sensitive information packets must be secure and protected from being hacked into
Nicole |  |  | Poor task scheduling from OpenAPS - wrong order algorithm, the dosing recommendation is always calculated before reading data and the recommendation is always made on the base of old data | 2 | Correct program flow must be specified and monitored
Nicole |  |  | Poor task scheduling from OpenAPS, - operating system, badly implemented call-flow | 2 | Correct scheduling must be specified and monitored
Nicole / Kate | Human - Human actor suddenly needs to remove the insulin pump and OpenAPS rig |  | Pump physically loses connection to body | 1 | Don't know - I assume this is defined in the user manual of the pump?
Nicole / Kate | Human - Human actor wants to change something with the OpenAPS rig / insulin pump and switches off the OpenAPS rig | RESUME HERE! | The battery and performance of an insulin pump is not tracked against the use and non-use of OpenAPS | 2 | If possible, OpenAPS must monitor insulin pump battery OpenAPS must make less than X requests to the insulin pump per minute
Nicole / Kate | Human - Human actor knows of a hazard which OpenAPS doesn't |  | The disk for RaspberryPi fills up and prevents the operating system from running and completing its processes | 1
Nicole / Kate | Human - Human actor accidently turns off or disconnects the OpenAPS rig (battery falls out, wires caught, human takes rig out of range of pump etc) |  | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the insulin pump (UCA-ROD-102: OpenAPS requests insufficient data from the insulin pump to make a reliable dosing recommendation) | 1 | OpenAPS algorithm must request all necessary/helpful information from insulin pump, regardless of what it is pulling from other components
Nicole / Kate | Human - Human actor mistinterprets a signal from the OpenAPS rig or pump or logs |  | The other information is (at least believed by OpenAPS to be) attainable from other components in the system so doesn't request all the data from the Nightscout | 1 | OpenAPS algorithm must request all necessary/helpful information from NightScout, regardless of what it is pulling from other components
Nicole / Kate | Human - Human actor believes the insulin dosing is wrong or not working (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) |  | The time that it takes for OpenAPS to make a dosing recommendation from receiving information has been lengthened, but the OpenAPS algorithm has not been adjusted for this | 2 | If the OpenAPS algorithm has a fixed time interval for making recommendations, this time interval should always be updated whenever there is a change in the algorithm's information intake interval or a delay in information intake or processing
| Human - OpenAPS is taking a while to begin and the human thinks this means it's not working properly |  | There is no feedback from the insulin pump to OpenAPS indicating that the insulin pump is not in a state to give a dosing recommendation | 1 | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user. OpenAPS must alert human actor if the insulin pump has stopped giving feedback OpenAPS should be able to interpret error signals from insulin pump and convey to the human actor
| Human - Human confusion (UCA-SO-100: Human actor provides Switch off (hold button goes into reset or another way) when the system is expected to be up and running) |  | Things stuck in the tube between the infusion set and the insulin pump | 1 | Must have watchdog timer or heartbeat communications between systems to detect if insulin pump is working properly and if not warn the user. OpenAPS must alert human actor if the insulin pump has stopped giving feedback OpenAPS should be able to interpret error signals from insulin pump and convey to the human actor
|  |  | This column is normally hidden | 1
| Human error (accident) (UCA-SO-1000: Human actor breaks the Power Off button (or another part of the OpenAPS rig)) |  | z N / A (not a Loss Scenario) - signal cannot reach insulin pump early (i.e. before we've sent it). Signal cannot be late (if it is sent the wireless signal will either reach the pump quickly or it will not reach the pump at all) | 1
| Human - Human does not see warning |  | z N/A, same applies to Nightscout as to the Insulin pump, too early or too late is umpossible | 1
|  |  | z N/A, see previous control action | 1
|  |  | z Same reasons as for above control action | 2
|  |  | #REF! | 42
0
0
0
| Human - Human is too busy or cannot be bothered to input data to the insulin pump |  |  | 0
| Human error (accident) UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime) |  |  | 0
| Human - Human has not read the instructions of the insulin pump properly / read wrong instructions |  |  | 0
| Human - Human has read insulin pump instructions but does not know how to input data into current insulin pump in use |  |  | 0
| Human - Human forgets to input data or that they can input data which is useful |  |  | 0
0
| Human - Human actor is experimenting |  |  | 0
| Human error (accident) UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing |  |  | 0
| Human - Human actor has not read the insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing) |  |  | 0
| Human - Human actor has not read or understood OpenAPS instructions for inputting data into the insulin pump |  |  | 0
| Human - Human actor has read old / deprecated instructions |  |  | 0
| Human - Instructions are complex and Human actor has forgetten them |  |  | 0
| Human error (accident) UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime |  |  | 0
| Human - Human trying to influence insulin dosing / glucose levels (UCA-UD-1000: Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime) |  |  | 0
| Human - Human does not read docs |  |  | 0
0
0
0
0
0
| Human - Human user wants their bolus dose rather than the basal recommendation (if they have additional information about their current state or upcoming activities) |  |  | 0
| Human error (accident) UCA-SB-100: Human actor sets a bolus (or basal) which unintentionally overrides OpenAPS temporary basals |  |  | 0
| Human - Human actor has not read or understood the OpenAPS instructions |  |  | 0
0
0
0
| Human - Human actor is experimenting |  |  | 0
Human error (accident) UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS
Human - Human actor believes the OpenAPS instructions are wrong
Human - Human actor has read old / deprecated instructions
Human - Human actor has not read or understood the OpenAPS instructions
Human - Human actor has forgetten the OpenAPS instructions (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS)
OpenAPS has received inaccurate glucose data from Nightscout
OpenAPS has gone out of range from Nightscout and not received updated glucose information
OpenAPS has gone out of range from the insulin pump and not received updated operating information
Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
OpenAPS clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
OpenAPS has received inaccurate glucose data from the human to override glucose data from Nightscout
The disk for RaspberryPi fills up and prevents the operating system from running and completing its processes
#REF!
#REF!
OpenAPS rig has differing timestamps to the insulin pump and believes the last recommendation was made recently
OpenAPS has received incorrect data from the human actor which conflicts with the rest of the current data and prevents a recommendation
Lack of algorithm programmer understanding of the information level required to make a dosing recommendation
OpenAPS rig has differing timestamps to the insulin pump and believes the last recommendation was made recently
OpenAPS has received inaccurate glucose data from Nightscout
Human actor has input too many carbs or gives incorrect information relating to exercise / other factors
OpenAPS is not informed of human bolus dosing
Insulin pump is configured with the wrong DIA / ISF / carb ratio settings
OpenAPS does not know that the human has missed a meal / unusual circumstances / travelling
OpenAPS has received inaccurate glucose data from Nightscout
Human actor has input too few carbs or gives incorrect information relating to exercise / other factors
Insulin pump is not operating correctly but OpenAPS is not informed of this
Insulin pump is configured with the wrong DIA / ISF / carb ratio settings
Human suddenly consumes a lot of food / sugar
OpenAPS may lack some information that conflicts with information that it has, which means it should hold back on making a decision
OpenAPS has not received correct data from the human actor which would conflict with the rest of the current data and prevent a recommendation
OpenAPS has received inaccurate glucose data from Nightscout
OpenAPS has received inaccurate glucose data from the human to override glucose data from Nightscout
OpenAPS does not pull all of the relevant insulin pump configuration information from the insulin pump
OpenAPS is having problems communicating with insulin pump
#REF!
#REF!
#REF!
#REF!
#REF!
There is no feedback from the insulin pump to OpenAPS indicating that the insulin pump is not in a state to give a dosing recommendation
Insulin pump is not aware that it is not in a state to make a dosing recommendation
Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
Insulin pump is not aware that it is not in a state to make a dosing recommendation
Insulin pump is not aware that it is not in a state to make a dosing recommendation
Insulin pump sends data to OpenAPS with the current timestamp rather than the time that the data was from
Insulin pump sends data without including any indicator that the data is not valid, recent data
Insulin pump clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
The battery and performance of an insulin pump is not tracked against the use and non-use of OpenAPS
Lack of algorithm programmer understanding of the information level required to make a dosing recommendation
Data is sent from the insulin pump which is actually null data or dummy data
OpenAPS is not informed of other processes in the schedule
CGM clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
Nightscout sends data to OpenAPS with the current timestamp rather than the time that the data was from
Lack of glucose data verification
OpenAPS has received inaccurate glucose data from Nightscout
CGM clock is set to the wrong time (changing timezone and recalibrating to local time, clocks changing, clock 'error' over time)
CGM sends data to OpenAPS with the current timestamp rather than the time that the data was from
The battery and performance of an insulin pump is not tracked against the use and non-use of OpenAPS
Lack of algorithm programmer understanding of the information level required to make a dosing recommendation
Data is sent from the insulin pump which is actually null data or dummy data
OpenAPS is not informed of other processes in the schedule
Human - Error in OpenAPS instructions (Wrong version number of the software)
Human - Incomplete instructions
Human - Flaw in the OpenAPS rig hardware setup design which the creators don't realise
Human - Docs out of date (UCA-PTR-100: Human actor sets up the OpenAPS rig wrong or insecurely)
Human - OpenAPS gives the wrong information to the human (possible wrong version number, possible multiple instances of same program etc.)
Human - Human actor cannot find information
Human - Error in OpenAPS instructions
Human - Incomplete instructions
Human - New insulin pump software has not been properly tested with OpenAPS
Human - OpenAPS gives the wrong information to the human (possible wrong version number, possible multiple instances of same program etc.)
Human - Docs out of date (when software has been updated)
Human - Human actor cannot find information
Human - Human actor cannot find information (installing dependencies)
Human - Docs out of date (Dependencies script out of date)
Human - Bug in dependencies script
Human - Vulnerability with one of the dependencies / software that OpenAPS people are unaware of
Human - Human is not notified when there are updates to OpenAPS are available
Human - Nobody at all is checking whether there are updates to OpenAPS system dependencies available
Human - Nobody monitors and tests the latest updates to determine if they are compatible with OpenAPS
Human - Human is not prompted to update system dependencies after the initial install (?)
Human - Error in OpenAPS instructions
Human - Incomplete instructions
Human - New OpenAPS software has not been properly tested with all the other devices (insulin pumps, CGM)
Human - Human actor cannot find information (working version)
Human - Docs out of date (working versions information)
Human - Human actor cannot find information (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS)
Human - Incomplete docs
Human - Docs out of date (UCA-S-101: Human actor installs unecessary additional system dependencies for OpenAPS)
Human - Human actor cannot find information (UCA-CO-1: There is a change in the OpenAPS loop setup, but the human actor does not update the configurations of OpenAPS accordingly)
Human - Incomplete or inaccurate docs, missing out a configuration setting
Human - There's a new potential device that hasn't been included in the configs
Human - Vulnerability to malicious actor interacting
Human - Human actor cannot find information (UCA-CO-100: Human actor selects the wrong option when configuring OpenAPS)
Human - Incomplete or inaccurate docs, missing out a configuration setting
Human - Vulnerability to malicious actor interacting
Human - Error in OpenAPS instructions
Human - Incomplete or inaccurate docs surrounding an algorithm
Human - Human actor doesn't get a notification email (UCA-CO-101: Human actor configures OpenAPS with an algorithm other than the latest oref0)
Human - Lack of feedback from OpenAPS (Human actor does not stop the loop when there is a problem with OpenAPS but OpenAPS is still issuing insulin doses)
Human - OpenAPS rig does not 'realise' it's not working well
Human - OpenAPS does not pull feedback from insulin pump (UCA-SO-1: Human actor does not stop the OpenAPS loop when there is a problem with OpenAPS / OpenAPS communication with other devices, but the OpenAPS loop is still issuing insulin doses)
Human - OpenAPS rig does not 'realise' the insulin pump is not dosing as instructed
Human - OpenAPS algorithm has a bug which makes it indicate that the OpenAPS system needs switching off when it doesn't
Human - Lack of feedback from rig or pump indicating that the OpenAPS looping is working (especially when dosing hasn't begun quite yet)
Human - There is no warning on the OpenAPS rig that OpenAPS is vulnerable to being broken
Human - Human does not see warning
Human - Error in insulin pump instructions (UCA-UD-1: Human actor does not provide data to the insulin pump which can improve insulin dosing recommendations (eg carb data at mealtime)
Human - Incomplete insulin pump instructions
Human - Insulin pump doesn't accomodate for the useful data
Human - Error in insulin pump instructions (UCA-UD-100: Human actor provides erroneous input of data affecting insulin dosing)
Human - Incomplete insulin pump or OpenAPS instructions
Human - Human actor is a different person to the user and doesn't know the accurate data of the user (parent doesn't know kid had some extra candy)
Human - Error in OpenAPS instructions
Human - Lack of information in docs (Human actor provides meal data to the insulin pump more than X minutes away from the actual mealtime)
Human - Error in OpenAPS instructions
Human - Incomplete OpenAPS instructions
Human - Another old set of OpenAPS instructions is available
Human - Human actor cannot find information
Human - OpenAPS instructions have an error (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS)
Human - OpenAPS instructions are incomplete (UCA-PM-100: Human actor configures insulin pump settings in a way that is not suitable for OpenAPS)
Human - OpenAPS does not make sure that the insulin pump settings are suitable for OpenAPS to operate
OpenAPS has gone out of range from the insulin pump and cannot issue a dosing recommendation OpenAPS has had wireless interference from the insulin pump and cannot issue a dosing recommendation
OpenAPS radio stick is not plugged into the OpenAPS rig properly
Environment has noisy radio frequency or electromagnetic field
OpenAPS radio frequency not tuned correctly
Equipment failiure
Insulin pump is blocked physically from the OpenAPS rig by another object
Extreme environmental changes
Explorer board not sitting properly on the Edison
#REF!
#REF!
#REF!
#REF!
#REF!
z Same reasons as for above control action
OpenAPS request is formatted erroneously
OpenAPS request is designed for a different style of insulin pump or different insulin pump software
z Same reasons as for above control action
OpenAPS request is formatted erroneously
OpenAPS request is designed for a different version of Nightscout
Insulin pump falls back to standard operation (not using OpenAPS)
#REF!
#REF!
#REF!
#REF!
#REF!
Insulin pump has received an operating data request from another party (UCA-ROD-1: OpenAPS does not get data from the insulin pump when it has not received data from the last X minutes )
Nightscout has received an operating data request from another party
z N / A (not a Loss Scenario) - signal cannot reach insulin pump early (i.e. before we've sent it). Signal cannot be late (if it is sent the wireless signal will either reach the pump quickly or it will not reach the pump at all)
#REF!
#REF!
#REF!
#REF!
#REF!
z N/A, see previous control action
z N/A, same applies to Nightscout as to the Insulin pump, too early or too late is umpossible
Insulin pump has run out of power
Insulin pump uses different units
Things stuck in the tube between the infusion set and the insulin pump
Pump physically loses connection to body
Insulin pump has run out of insulin
#REF!
#REF!
#REF!
#REF!
#REF!
Insulin pump has run out of power
Insulin software / firmware has been updated and insulin pump now lacks compatibiliity with OpenAPS
Nightscout is not online
Nightscout / CGM software / firmware has been updated and it now lacks compatibiliity with OpenAPS
Insulin pump falls back to standard operation (not using OpenAPS)
#REF!
#REF!
#REF!
#REF!
#REF!
Liquid in the tubing connector has caused the infusion set to get blocked
Insulin pump scheduling causes commands from OpenAPS to be executed too late
#REF!
#REF!
#REF!
#REF!
#REF!

## Level 2 UCAs
Controller | Control Action | Not Providing | NP Controller Constraints | Providing | P Controller Constraints | Timing | T Controller Constraints | Duration | D Controller Constraints | Notes
Raspi OS | RSA: Set up algorithm | UCA-RSA-1: Raspi OS does not update the OpenAPS configurations (or prompt the human operator to do so) when there is a change in the OpenAPS loop setup that should affect the configurations | CC-RSA-1: When OpenAPS configurations are/become incorrect, this should be detected by or told to the RasPi which should then notify the human operator and if the configuration is dangerous OpenAPS should update the configuration itself and notify the human operator of the new value | UCA-RSA-100: Raspi OS selects a different option to that which the human operator entered or intended when configuring OpenAPS | CC-RSA-100: When the human operator enters or intends an option when configuring OpenAPS, the Raspi OS must select that option or inform the human operator that their configuration is unsafe and what the new configuration is | UCA-RSA-1000: Raspi OS updates the algorithm or configurations too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings | CC-RSA-1000: If the configuration files for the algorithm have changed but the algorithm is still running with the old configurations, OpenAPS should notify the human operator
|  | UCA-RSA-2: Raspi OS does not successfully set up the algorithm when the human operator runs the script to set up the algorithm | CC-RSA-2: When the human operator correctly runs the script to set up the algorithm, the RasPi OS must make OpenAPS install and start looping the algorithm | UCA-RSA-101: Raspi OS configures OpenAPS with a nonoptimal algorithm instead of oref0 | CC-RSA-101: Raspi OS should warn the human operator that OpenAPS is about to be configured with a non-recommended algorithm before proceeding with the setup
| STAO: Start OpenAPS | UCA-STAO-1: Raspi OS does not start OpenAPS loop when Raspi launches and there is an OpenAPS loop set up | CC-STAO-1: If there is a loop set up, when the Raspi OS is turned on it should start the OpenAPS loop | UCA-STAO-100: Raspi OS starts a different application and not OpenAPS | CC-STAO-100: OpenAPS should not cause the RasPi to launch a different application | UCA-STAO-1000: Raspi OS takes more than X seconds to start the OpenAPS loop | CC-STAO-1000: Raspi OS should take less than X seconds to start the OpenAPS loop
|  | UCA-STAO-2: Raspi OS does not start OpenAPS when the OpenAPS loop is installed | CC-STAO-2: When the OpenAPS loop is set up, Raspi OS should then start the OpenAPS loop | UCA-STAO-101: Raspi OS starts an older version of OpenAPS | CC-STAO-101: When it is turned on, Raspi OS should launch the latest working version of OpenAPS available
|  | UCA-STAO-3: Raspi OS does not start OpenAPS after there has been a completed update to OpenAPS | CC-STAO-3: After OpenAPS has finished an update and is ready to continue, the Raspi OS should start the OpenAPS loop again or inform the human operator | UCA-STAO-102: Raspi OS starts OpenAPS twice | CC-STAO-102: By default Raspi OS should just be running OpenAPS once whenever it is turned on
|  |  |  | UCA-STAO-103: Raspi OS starts OpenAPS loop when OpenAPS is not working properly | CC-STAO-103: If OpenAPS is generating hazardous glucose levels and still controlling the insulin pump, the Raspi should stop the OpenAPS loop
| RPD: Provide data from other devices to OpenAPS | UCA-RPD-1: Raspi OS does not relay data from other devices to OpenAPS app when OpenAPS is running | CC-RPD-1: Raspi OS must relay data from other devices to OpenAPS app when OpenAPS is running. CC-RPD-2: If OpenAPS has not received any data from Raspi OS in X minutes, OpenAPS should stop controlling the insulin pump. | UCA-RPD-100: Raspi OS relays data from other devices to OpenAPS app inaccurately | CC-RPD-100: Raspi OS must relay data from other devices to OpenAPS app accurately | UCA-RPD-1000: Raspi OS takes more than X seconds to relay data from other devices to OpenAPS app | CC-RPD-1000: Raspi OS must take less than X seconds to relay data from other devices to OpenAPS app
|  |  |  | UCA-RPD-101: Raspi OS relays data from other devices to OpenAPS app when other devices have not sent data to OpenAPS | CC-RPD-101: Raspi OS must not relay data from other devices to OpenAPS app that is not actually from the other devices
| STOO: Stop OpenAPS | UCA-STOO-1: Raspi OS does not stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes | CC-STOO-1: Raspi OS must stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes | UCA-STOO-100: Raspi OS stops OpenAPS when OpenAPS is running effectively and there has been no instruction from the human operator to stop OpenAPS | CC-STOO-100: Raspi OS should not stop OpenAPS when OpenAPS is running effectively and there has been no instruction from the human operator to stop OpenAPS | UCA-STOO-1000: Raspi OS takes more than X minutes to stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes | CC-STOO-1000: Raspi OS should take less than X minutes to stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes
| SB: Set Bolus | UCA-SB-1: OpenAPS does not set a supermicrobolus for a meal when it is running the oref1 algorithm |  | UCA-SB-100 : OpenAPS algorithm recommends a too high bolus dosing recommendation based on current glucose levels, active insulin and human insulin pump input |  | UCA-SB-1000 : OpenAPS provides a bolus recommendation too late after a meal time
|  |  |  | UCA-SB-101 : OpenAPS algorithm recommends a too low bolus dosing recommendation based on current glucose levels, active insulin and human insulin pump input |  | UCA-SB-1001 : OpenAPS provides a bolus recommendation too soon after the last bolus dose from the insulin pump
UCA-SB-102: OpenAPS makes a bolus recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
UCA-SB-103 : OpenAPS makes a bolus dosing recommendation which overrides the human’s safe requested level
UCA-SB-104 : OpenAPS provides a bolus dosing rate when the insulin pump settings are not configurated correctly (*1)
UCA-SB-105: OpenAPS makes a bolus recommendation when it is running an algorithm that should only set temporary basal rates
UCA-SB-106 : OpenAPS makes bolus recommendations too frequently (repetitive stuck)
UCA-STB-107 : OpenAPS makes a bolus recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
| SI: Send instruction to Insulin pump | UCA-SI-1: The RasPi OS does not send a command to the insulin pump when the OpenAPS app has requested this | CC-SI-1: The RasPi OS should send a command to the insulin pump when the OpenAPS app has requested this | UCA-SI-100: The RasPi OS sends a instruction to the Insulin Pump when it has not been sent a command by the OpenAPS app or a prompt from the recipient | CC-SI-100: The RasPi OS should not send an instruction to the Insulin Pump when it has not been sent a command by the OpenAPS app or a prompt from the recipient | UCA-SI-1000: The RasPi OS sends an instruction to Insulin pump more than X seconds after the command has been invoked by OpenAPS | CC-SI-1000: The RasPi OS should send instructions to Insulin pump less than X seconds after the command has been invoked by OpenAPS
|  |  |  | UCA-SI-101: The RasPi OS sends the wrong command to the Insulin Pump | CC-SI-101: The RasPi OS should always send the same command that OpenAPS informed it to to the Insulin Pump
|  |  |  | UCA-SI-102: The RasPi OS sends a command to the Insulin Pump when the recipient is not able to receive or action the command | CC-SI-102: When the RasPi OS sends a command to the Insulin Pump that the Insulin pump is not able to receive or action, the human operator should be informed and if safety critical the insulin pump should stop using the OpenAPS dosing recommendations
|  |  |  | UCA-SI-103: The RasPi OS sends a command to the Insulin Pump based on inaccurate information | CC-SI-103: The RasPi OS should send all commands to the Insulin Pump based on accurate information and calculations
|  |  |  | UCA-SI-104: The RasPi OS sends an unsucessful instruction or data request to the Insulin Pump | CC-SI-104: When the RasPi OS sends an unsucessful instruction or data request to the Insulin Pump, the human operator should be informed and if safety critical OpenAPS should stop controlling the insulin pump.
| SN: Send instruction to NightScout | UCA-SN-1: The RasPi OS does not send a command to NightScout when the OpenAPS app has requested this | CC-SN-1: The RasPi OS should send a command to the NightScout when the OpenAPS app has requested this | UCA-SN-100: The RasPi OS sends a instruction to the NightScout when it has not been sent a command by the OpenAPS app or a prompt from the recipient | CC-SN-100: The RasPi OS should not send an instruction to NightScout when it has not been sent a command by the OpenAPS app or a prompt from the recipient | UCA-SN-1000: The RasPi OS sends an instruction to NightScout more than X seconds after the command has been invoked by OpenAPS | CC-SN-1000: The RasPi OS should send instructions to NightScout less than X seconds after the command has been invoked by OpenAPS
|  |  |  | UCA-SN-101: The RasPi OS sends the wrong command to the NightScout | CC-SN-101: The RasPi OS should always send the same command that OpenAPS informed it to to NightScout
|  |  |  | UCA-SN-102: The RasPi OS sends a command to the NightScout when the recipient is not able to receive or action the command | CC-SN-102: When the RasPi OS sends a command to the NightScout that NightScout is not able to receive or action, the human operator should be informed and if safety critical the insulin pump should stop using the OpenAPS dosing recommendations
|  |  |  | UCA-SN-103: The RasPi OS sends a command to NightScout based on inaccurate information | CC-SN-103: The RasPi OS should send all commands to NightScout based on accurate information and calculations
|  |  |  | UCA-SN-104: The RasPi OS sends an unsucessful instruction or data request to NightScout | CC-SN-104: When the RasPi OS sends an unsucessful instruction or data request to the NightScout, the human operator should be informed and if safety critical OpenAPS should stop controlling the insulin pump.
OpenAPS app | OSA: Set up algorithm and loop | UCA-OSA-1: OpenAPS app does not update the OpenAPS configurations (or prompt the human operator to do so) when there is a change in the OpenAPS loop setup that should affect the configurations | CC-OSA-1: When OpenAPS configurations are/become incorrect, this should be detected by OpenAPS which should then notify the human operator and if the configuration is dangerous the OpenAPS rig should update the configuration itself and notify the human operator of the new value | UCA-OSA-100: OpenAPS app selects a different option to that which the human operator entered or intended when configuring OpenAPS | CC-OSA-100: When the human operator enters or intends an option when configuring OpenAPS, OpenAPS must select that option or inform the human operator that their configuration is unsafe and what the new configuration is | UCA-OSA-1000: OpenAPS app updates the algorithm or configurations too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings | CC-OSA-1000: If the configuration files for the algorithm have changed but the algorithm is still running with the old configurations, OpenAPS should notify the human operator
|  | UCA-OSA-2: OpenAPS app does not successfully set up the algorithm when the script is run to set up the algorithm | CC-OSA-2: When the human operator correctly runs the script to set up the algorithm, OpenAPS must install and start looping the algorithm | UCA-OSA-101: OpenAPS sets up an unsafe algorithm | CC-OSA-101: OpenAPS should warn the human operator that OpenAPS is about to be configured with a non-recommended algorithm before proceeding with the setup
| OPD: Provide data from other devices to algorithm | UCA-OPD-1: OpenAPS app does not relay relevant data from other devices to the algorithm | CC-OPD-1: OpenAPS app must relay data from other devices to algorithm when OpenAPS is running. CC-OPD-2: If algorithm has not received any data from OpenAPS app in X minutes, the algorithm should instruct OpenAPS to stop controlling the insulin pump. | UCA-OPD-100: OpenAPS app relays data from other devices to algorithm inaccurately | CC-OPD-100: OpenAPS app must relay data from other devices to the algorithm accurately | UCA-OPD-1000: OpenAPS app takes more than X seconds to relay data from other devices to the algorithm | CC-OPD-1000: OpenAPS app must take less than X seconds to relay data from other devices to the algorithm
|  |  |  | UCA-OPD-101: OpenAPS app relays data regarding other devices to algorithm when other devices have not sent data to OpenAPS | CC-OPD-101: OpenAPS app must not relay data from other devices to the algorithm that is not actually from the other devices
| STOA: Stop algorithm | UCA-STOA-1: OpenAPS app does not stop algorithm and fall back to regular pump operation when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes | CC-STOA-1: OpenAPS app must stop algorithm and fall back to regular pump operation when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes | UCA-STOA-100: OpenAPS app stops algorithm when OpenAPS is giving good dosing recommendations to the insulin pump and the system is running correctly and there has been no instruction from the human operator to stop OpenAPS | CC-STOA-100: OpenAPS app should not stop the algorithm when OpenAPS is giving good dosing recommendations to the insulin pump and the system is running correctly and there has been no instruction from the human operator to stop OpenAPS | UCA-STOA-1000: OpenAPS app takes more than X minutes stop algorithm when the algorithm is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes | CC-STOA-1000: OpenAPS app should take less than X minutes to stop algorithm when the algorithm is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes
Infrastructure developer | IDEN: Email notification (to human operator) | UCA-IDEN-1: InfDev does not send an e-mail and the HuOp does not know that the OpenAPS app has been updated | CC-IDEN-1: When the OpenAPS app has been updated, the InfDev must send a notification to inform the HuOp | UCA-IDEN-100: InfDev sends out an email about something that is not there or an update that has not happened | CC-IDEN-100: The InfDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this | UCA-IDEN-1000: InfDev sends out an email with safety critical information X hours too late | CC-IDEN-1000: InfDev should send out an email with safety critical information within X hours of discovering this information | UCA-IDEN-10000: InfDev sends out an email with incomplete information (so it seems everything is ok, but actually some important information is missing) | CC-IDEN-10000: The InfDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-IDEN-100)
|  | UCA-IDEN-2: InfDev does not send an email, when there should be one to communicate safety critical information | CC-IDEN-2: When there is safety critical information related to the OpenAPS app, the InfDev must send a notification to inform the HuOp | UCA-IDEN-101: InfDev sends out an email with wrong information | CC-IDEN-101: The InfDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-IDEN-100) | UCA-IDEN-1001: InfDev sends out an email too early, before the implementation/change is actually available | CC-IDEN-1001: If the InfDev is sending out an email before the implementation/change is actually available, this should be mentioned in the e-mail | UCA-IDEN-10001: InfDev sends out an email that contains both valid information and information about features/changes/implementations that are not yet there. | CC-IDEN-10001: The InfDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-IDEN-100)
|  |  |  | UCA-IDEN-102: InfDev sends a duplicate e-mail message to the HuOp that the OpenAPS app has been updated, causing the HuOp to be confused | CC-IDEN-102: The InfDev should not send a duplicate e-mail message to the HuOp that the OpenAPS app has been updated (it is very unlikely that they will)
|  |  |  | UCA-IDEN-103: InfDev sends out an email with missing information | CC-IDEN-103: Notification e-mails from the InfDev should contain all helpful information which is not basic knowledge, or links to this information
|  |  |  | UCA-IDEN-104: InfDev sends out an email in the wrong language | CC-IDEN-104: Notification e-mails from the InfDev should always be in the default language of English, unless the addressee has requested otherwise
| INFO: Forum (to human operator) | UCA-INFO-1: InfDev does not provide a forum and the HuOp is not aware of any concerns or known issues for the Algorithm or able to contact the InfDev | CC-INFO-1: InfDev must provide a forum for the HuOps to comment on | UCA-INFO-100: InfDev provides erroneous information on the forum | CC-INFO-100: InfDev should not provide erroneous information on the forum and if they do then they should update this when they become aware of this | UCA-INFO-1000: InfDev updates the forum to reflect an OpenAPS app change that is not available yet | CC-INFO-1000: If the InfDev is updating the forum before the implementation/change is actually available, this should be mentioned in the forum update
|  | UCA-INFO-2: InfDev fails to update the forum and the HuOp is not aware of a new issue found for the algorithm | CC-INFO-2: InfDev must update the forum with urgent information (all other important information can be communicated through e-mail) | UCA-INFO-101: InfDev provides an insufficient answer | CC-INFO-101: InfDev should make sure that they provide all necessary information in their answers to forum questions and the HuOp should have means to communicate if they are not satisfied with the answer | UCA-INFO-1001: InfDev updates the forum well after the OpenAPS app has been updated | CC-INFO-1001: InfDev should update the forum with safety critical information within X hours of discovering this information
|  | UCA-INFO-3: InfDev does not provide an answer to a question from the HuOp | CC-INFO-3: InfDev must provide an answer to all unanswered questions from the HuOp on the forum | UCA-INFO-102: InfDev provides their answer to the wrong question | CC-INFO-102: It should be clear to which question the answer from the InfDev is for
|  | UCA-INFO-4: InfDev does not update an answer when it becomes false | CC-INFO-4: InfDev should update answers to common FAQs within X hours when they become false (if there is an FAQs section)
| UTOO: Update (to OpenAPS app) | UCA-UTOO-1: InfDev fails to update the OpenAPS app to address a known issue | CC-UTOO-1: InfDev should update the OpenAPS app to address a known issue or instruct the HuOp to do this | UCA-UTOO-100: InfDev provides an erroneous OpenAPS app update, e.g. installs an older version of the OpenAPS app | CC-UTOO-100: The InfDev should test any OpenAPS app update and make sure they provide the correct tested version to the HuOps | UCA-UTOO-1000: InfDev is slow in updating the OpenAPS app | CC-UTOO-1000: When the OpenAPS app has a safety-critical bug, the InfDev should update the OpenAPS app to fix this within X hours
Algorithm developer | ALEN: E-mail notification (to human operator) | UCA-ALEN-1: AlgDev does not send an e-mail and the HuOp does not know that the algorithm has been updated | CC-ALEN-1: When the algorithm has been updated, the AlgDev must send a notification to the HuOp | UCA-ALEN-100: AlgDev sends out an email about something that is not there or an update that has not happened | CC-ALEN-100: The AlgDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this | UCA-ALEN-1000: AlgDev sends an e-mail well after the algorithm has been updated, unbeknownst to the HuOp (similar to other UCAs) | CC-ALEN-1000: AlgDev should send out an email with safety critical information within X hours of discovering this information | UCA-ALEN-10000: AlgDev sends out an email with incomplete information (so it seems everything is ok, but actually some important information is missing) | CC-ALEN-10000: The AlgDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-ALEN-100)
|  | UCA-ALEN-2: AlgDev does not send an e-mail to communicate safety critical information | CC-ALEN-2: When there is safety critical information related to the algorithm, the AlgDev must send a notification to inform the HuOp | UCA-ALEN-101: AlgDev sends out an email with wrong information | CC-ALEN-101: The AlgDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-ALEN-100) | UCA-ALEN-1001: AlgDev sends an e-mail before the algorithm has been updated and the HuOp thinks the algorithm has been updated when it hasn't (similar to other UCAs) |  | UCA-ALEN-10001: AlgDev sends out an email that contains both valid information and information about features/changes/implementations that are not yet there. | CC-ALEN-10001: The AlgDev should not send out a false e-mail and if they do then they should send out another notification when they become aware of this (same as CC-ALEN-100)
|  |  |  | UCA-ALEN-102: AlgDev sends a duplicate e-mail message to the HuOp that the algorithm has been updated, causing the HuOp to be confused | CC-ALEN-102: The AlgDev should not send a duplicate e-mail message to the HuOp that the algorithm has been updated (it is very unlikely that they will)
|  |  |  | UCA-ALEN-103: AlgDev sends out an email with missing information | CC-ALEN-103: Notification e-mails from the AlgDev should contain all helpful information which is not basic knowledge, or links to this information
|  |  |  | UCA-ALEN-104: AlgDev sends out an email in the wrong language | CC-ALEN-104: Notification e-mails from the AlgDev should always be in the default language of English, unless the addressee has requested otherwise
| ALFO: Forum (to human operator) | UCA-ALFO-1: AlgDev does not provide a forum and the HuOp is not aware of any concerns or known issues for the Algorithm or able to contact the AlgDev | CC-ALFO-1: AlgDev must provide a forum for the HuOps to comment on | UCA-ALFO-100: AlgDev provides erroneous information on the forum, misleading the HuOp | CC-ALFO-100: AlgDev should not provide erroneous information on the forum and if they do then they should update this when they become aware of this | UCA-ALFO-1000: AlgDev updates the forum to reflect an algorithm that is not available yet | CC-ALFO-1000: If the AlgDev is updating the forum before the implementation/change is actually available, this should be mentioned in the forum update
|  | UCA-ALFO-2: AlgDev fails to update the forum and the HuOp is not aware of a new issue found for the algorithm | CC-ALFO-2: AlgDev must update the forum with urgent information (all other important information can be communicated through e-mail) | UCA-ALFO-101: AlgDev provides an insufficient answer | CC-ALFO-101: AlgDev should make sure that they provide all necessary information in their answers to forum questions and the HuOp should have means to communicate if they are not satisfied with the answer | UCA-ALFO-1001: AlgDev updates the forum well after the algorithm has been updated | CC-ALFO-1001: AlgDev should update the forum with safety critical information within X hours of discovering this information
|  | UCA-ALFO-3: AlgDev does not provide an answer to a question from the HuOp | CC-ALFO-3: AlgDev must provide an answer to all unanswered questions from the HuOp on the forum | UCA-ALFO-102: AlgDev provides their answer to the wrong question | CC-ALFO-102: It should be clear to which question the answer from the AlgDev is for
|  | UCA-ALFO-4: AlgDev does not update an answer when it becomes false | CC-ALFO-4: AlgDev should update answers to common FAQs within X hours when they become false (if there is an FAQs section)
| ALUP: Update (to algorithm) | UCA-ALUP-1: AlgDev fails to update the algorithm to address a known issue | CC-ALUP-1: AlgDev should update the algorithm to address a known issue or instruct the HuOp to do this | UCA-ALUP-100: AlgDev provides an erroneous algorithm update, e.g. installs an older version of the algorithm | CC-ALUP-100: The AlgDev should test any algorithm update and make sure they provide the correct tested version to the HuOps | UCA-ALUP-1000: AlgDev is slow in updating the algorithm | CC-ALUP-1000: When the algorithm has a safety-critical bug, the AlgDev should update the algorithm to fix this within X hours

## Level 2 Loss Scenario Analysis
Controller | Control Action | UCA | True Statement | Unsafe Controller Behaviour Type (Information given to controller is correct) | Controller Belief | Loss Scenario Type 1 | Requirement 1 | Inadequate Information and Feedback Type (information given to controller is incorrect) | Controller Belief | Loss Scenario Type 2 | Requirement 2 | Control Path Scenario - | How could this happen? | Requirement | Control Path Scenario - | How could this happen? | Requirement | Control Path Scenario - | How could this happen? | Requirement | Controlled Process Scenario | How could this happen? | Requirement | Controlled Process Scenario | How could this happen? | Requirement | Controlled Process Scenario | How could this happen? | Requirement
Raspberry Pi User Iterface | Install dependencies from Docs and Dependencies (ID) | UCA-ID-1: User Interface does not provide install dependencies when Human Actor provides valid OpenAPS configure command | Human actor provides valid OpenAPS configure command | Raspbian (RasPi OS) does not provide install dependencies though the information given to RasPi OS indicates that the OpenAPS configure command is valid | User Interface believes OpenAPS configure command is valid | System shutdown or interruption before/when installing dependencies | OpenAPS should be able to detect if dependencies are completely installed to match the current OpenAPS algorithm | Human actor provides valid OpenAPS configure command but the information given to RasPi OS does not indicate this |  | RasPi OS bug | Raspberry OS is up to date
[6 empty] | Internet cuts out before/when installing dependencies | (See above)
|  | UCA-ID-2: User Interface does not install all of the necessary dependencies when Human Actor provides valid OpenAPS configure command | Human actor provides valid OpenAPS configure command | Raspbian (RasPi OS) does not install all of the necessary dependencies though the information given to RasPi OS indicates that the OpenAPS configure command is valid |  | Some dependencies are not compatible with the OS version that the Raspberry Pi is running | OpenAPS should be able to detect if dependencies are completely installed to match the current OpenAPS algorithm | Human actor provides valid OpenAPS configure command with all of the necessary dependencies but the information given to RasPi OS does not indicate this |  | Human should make sure that the script has been run successfully Human should make sure OpenAPS installation has been successful OpenAPS should have a script which can be run to detect if all dependencies are completely installed | Human should make sure that the script has been run successfully Human should make sure OpenAPS installation has been successful OpenAPS should have a script which can be run to detect if all dependencies are completely installed
Lack of memory space on the Raspberry Pi causes some dependencies not to be install
A misspelt / non-existent dependency or incorrect install command in the OpenAPS setup script causes some dependencies not to be installed
[6 empty] | Whole dependencies setup script could have an error / be missing a necessary line or command | OpenAPS installation script must be reviewed There should be a script to test that all the required dependencies have been installed which is seperate to the installation script OpenAPS install script should have pre-install and post-install scripts, pre-install makes sure that OpenAPS has dependencies before the installation of OpenAPS, post install makes sure that everything which is required has been installed
Script is run in or installs dependencies in the wrong place
|  | UCA-ID-100: User Interface makes a software update to the OpenAPS dependencies, which makes it incompatible with other device(s) in the system (eg pump, CGM) | Software update makes OpenAPS incompatible with other device(s) in the system | User Interface makes a software update to the OpenAPS dependencies, though information indicates that makes it incompatible with other device(s) in the system (eg pump, CGM) | User interface believes the update makes OpenAPS incompatible with other device(s) in the system | OpenAPS dependencies are outdated or not compatible with other system components after new distro package update or upgrade | OpenAPS should check whether the software update is effective and conforming to OpenAPS recommendatoins in the context of the versions of the other components in the users system | Software update makes OpenAPS incompatible with other device(s) in the system but OpenAPS information does not indicate this |  | Different version of any dependency needed for other RasPi OS apps or the OpenAPS app | OpenAPS should require specific versions of dependencies which have been tested to be compatible with all of the components
[6 empty] | RasPi OS bug | Covered |  |  | RasPi OS does not know the dependencies of other compnents of the system
|  | UCA-ID-101: User Interface installs unecessary additional dependencies for OpenAPS | Additional dependencies are unecessary for OpenAPS | User Interface installs additional dependencies that it believes are unecessary for OpenAPS | User Interface believes that additional dependencies are not necessary for OpenAPS | Human install additional dependencies and Raspbian doesn't stop this | Human should be warned in documentation not to install too many packages | The additional dependencies are unecessary for OpenAPS but the Linux UI information does not indicate this |  | OpenAPS installation script has a bug and includes wrong dependencies
Raspbian installs unecessary additional dependencies when installing the necessary dependencies
[6 empty] | Raspbian is updating OS version and installs additional dependencies (which could conflict with OpenAPS dependencies) | There should be continuous testing of OpenAPS dependencies with updated OS
A package in Raspbian has been hacked into
The RasPi itself has been hacked into
|  | UCA-ID-102: User Interface makes a software update to the OpenAPS rig when OpenAPS is in the middle of communicating with another device or doing something that should not be interrupted | OpenAPS is in the middle of doing something that should not be interrupted | User Interface makes a software update to the OpenAPS rig when it believes OpenAPS should not be interrupted | User Interface believes that OpenAPS should not be interrupted | User Interface does not account for OpenAPS state correctly when making software updates | When there are software updates on the RasPi, OpenAPS should stop making recommendations | OpenAPS is in the middle of doing something that should not be interrupted but the Linux UI information does not indicate this |  | OpenAPS does not indicate to the RasPi when it is in a process that should not be interrupted
[6 empty] | Human makes an OpenAPS software update when the OpenAPS system is running and the RasPi does not stop or warn against this | Same as above
|  | UCA-ID-103: User Interface removes necessary dependencies for OpenAPS | The dependencies are necessary for OpenAPS | User Interface removes dependencies that it believes are necessary for OpenAPS | User Interface believes the dependencies are necessary for OpenAPS | Human removes the dependencies that are necessary for OpenAPS | Covered | The dependencies are necessary for OpenAPS but the Linux UI information does not indicate this |  | OpenAPS dependencies setup script is not updated
[6 empty] | RasPi has been hacked into | Human user should be warned of any possible security threats
[6 empty] | RasPi software update removes OpenAPS dependencies | Covered
RasPi bug
Dependency removes itself
OpenAPS dependency is updated and has changed name (so now has different name to OpenAPS dependency)
|  | UCA-ID-1000: User Interface installs dependencies after OpenAPS has been configured and started looping | OpenAPS has been configured and started looping | User Interface installs dependencies when it believes OpenAPS has already been configured and started looping | User Interface believes OpenAPS has already been configured and started looping | A bug in OpenAPS loop means that it installs dependencies as part of the loop | Out of scope | OpenAPS has been configured and started looping but the Linux UI information does not indicate this |  | -
[6 empty] | Dependencies are installed on the Raspberry Pi independently of the OpenAPS loop | Out of scope
|  | UCA-ID-1001: User Interface installs dependencies before Raspbian dependencies have been updated to be able to use the new versions | Raspbian dependencies are not yet compatible with new OpenAPS dependencies | User Interface installs dependencies when it believes Raspbian dependencies has not been updated to be able to use the new versions | User Interface believes Raspbian deps have not been updated to be able to use the new versions |  |  | Raspbian dependencies are not yet compatible with new OpenAPS dependencies but the Linux UI does not indicate this |  | -
Raspi OS | Set up algorithm (RSA) | UCA-RSA-1: Raspi OS does not update the OpenAPS configurations (or prompt the human operator to do so) when there is a change in the OpenAPS loop setup that should affect the configurations | There has been a change in the OpenAPS loop setup that should affect the configurations | Raspi OS does not update the OpenAPS configurations (or prompt the human operator to do so) when it believes there is a change in the OpenAPS loop setup that should affect the configurations | User Interface believes there has been a change in the OpenAPS loop setup that should affect the configurations | There is no information in the code/configuration that allows for automated compatibility checks | OpenAPS app shall remind the user to update configs when there has been a change in devices available OpenAPS app shall provide information to the user on how to update relevant configs when the devices available have changed OpenAPS documentation should provide information about how to set configs based on devices available | There has been a change in the OpenAPS loop setup that should affect the configurations but the Linux UI info does not indicate this |  | If there is information in the code/configuration that allows compatibility checks, the version information is not updated, so that the compatiblity algorithm thinks everything is still on its old configuration | OpenAPS lib setup shall be updated to be kept in sync with available devices
[6 empty] | If there is information in the code/configuration that allows compatibility checks, the compatibility checks are not implemented or errorneously implemented |  |  |  | There have been additional configurations added in the latest OpenAPS version so the version that is being used needs to be updated | OpenAPS shall notify the user when there are additional configurations available to be set
[6 empty] | The interface detects the configuration mismatch, but does not show this on the UI | OpenAPS app shall display accurate current status of connected devices to user
OpenAPS is not set up to prompt the human to update configs when it detects a change that should be reflected in configs or detects inaccurate configs
|  | UCA-RSA-2: Raspi OS does not successfully set up the algorithm when the human operator runs the script to set up the algorithm | Human operator has run the script to set up the algorithm | Raspi OS does not successfully set up the algorithm when it believes that the human operator runs the script to set up the algorithm |  | Dependencies on the RaspiOS are not / cannot be the correct version to install the algorithm correctly | For each RasPI OS dependency, the version(s) that are compatible each supported version of OpenAPS must be documented | Human operator has run the script to set up the algorithm but Raspi OS statuses do not indicate successful installation |  | Installation script aborts or crashes partway through | Human operator must have access to the command line so they can see whether the installation script suceeded Algorithm installation script should be tested on all versions of Raspbian OS that expected to be supported and the results should be feed into the documentation supported versions
[6 empty] | Raspi OS version is not supported by OpenAPS | For each RasPI OS dependency, the version(s) that are compatible each supported version of OpenAPS must be documented
[6 empty] | Raspi OS has packaging problems / broken packages which are causing problems installing dependencies | RasPi OS must not have broken dependencies and packages
|  | UCA-RSA-100: Raspi OS selects a different option to that which the human operator entered or intended when configuring OpenAPS | Human operator entered or intended a different configuration option than is selected | Raspi OS selects a different option to that it believes the human operator entered or intended when configuring OpenAPS | User Interface believes the human operator entered or intended a different option | Raspi OS bug or malfunction | OpenAPS developers shall validate that the OpenAPS API work with the version of Raspi OS and project supported hardware | Human operator entered or intended a different configuration option than is selected but the Linux UI information does not indicate this |  | Issue with recognizing the operator input | OpenAPS should allow the user to select the keyboard that they are using
[6 empty] | selection of a different memory section than the one the information entered by the operator is stored at (memory address issue) | Covered by above
[6 empty] | issue with passing the entered information on to beyond the user interface | Covered by above
|  | UCA-RSA-101: Raspi OS configures OpenAPS with a nonoptimal algorithm instead of oref0 | Algorithm is not optimal for OpenAPS | Raspi OS configures OpenAPS with an algorithm that it believes is nonoptimal instead of oref0 | User Interface believes algorithm is non-optimal | Raspi OS bug or malfunction | Covered by UCA-OSA-101 (possibly delete this UCA?) | Algorithm is not optimal for OpenAPS but the Linux UI information does not indicate this |  | Algorithm has safety flaws that noone knows about
[6 empty] | Human confirms that the algorithm is safe when it has not been tested properly |  |  |  | Linux UI misidentifies the algorithm as a recommended algorithm
|  | UCA-RSA-1000: Raspi OS updates the algorithm or configurations too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings | OpenAPS has already started to operate with the current version | User Interface updates too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings, when it believes that OpenAPS has already started with previous algorithm settings | User Interface believes that it's too late for the algorithm to encorporate the update | Raspi OS bug or malfunction | Requirements here Covered by algorithm delay list | OpenAPS has already started to operate with the current version but the Linux UI information does not indicate this |  | Issue with scheduling the tasks | Requirements here Covered by algorithm delay list
[6 empty] | Some other action is blocking the update routine | Requirements here Covered by algorithm delay list
[6 empty] | Scheduling/Interrupt issue | Requirements here Covered by algorithm delay list
[6 empty] | The Human operator starts running the algorithm before reconfiguring which algorithm to run | Requirements here Covered by algorithm delay list
| Start OpenAPS (STAO) | UCA-STAO-1: Raspi OS does not start OpenAPS loop on launch when there is an OpenAPS loop set up | RasPi OS launces and there is an OpenAPS loop set up | Raspi OS does not start the OpenAPS loop on launch when it believes there is an OpenAPS loop set up |  | Bug in OpenAPS code | Covered by Controller Constraint, CC-STAO-1: If there is a loop set up, when the Raspi OS is turned on it should start the OpenAPS loop | Raspi OS does not recognise there is a OpenAPS loop set up |  | OpenAPS loop is set up in the wrong location on RasPi OS
[6 empty] | Incompatible OpenAPS and RasPi versions | OpenAPS developers shall specify the versions of RaspiOS which are support on each version of OpenAPS OpenAPS app shall do a compatibility check on whether it is compatible with the current RaspiOS version |  |  | OpenAPS files have been corrupted
[6 empty] | Instruction to start OpenAPS has been cleared from Raspi OS boot instructions | On boot, RaspiOS shall automatically start OpenAPS
[6 empty] | Raspi OS was booted abnormally | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
[6 empty] | Raspi OS crash | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
|  | UCA-STAO-2: Raspi OS does not start OpenAPS when the OpenAPS loop is installed | OpenAPS loop is installed | Raspi OS does not start OpenAPS when it believes the OpenAPS loop is installed |  | Bug in OpenAPS code | Covered by Controller Constraint, CC-STAO-1: If there is a loop set up, when the Raspi OS is turned on it should start the OpenAPS loop | Raspi OS does not recognise that the OpenAPS loop has been installed |  | OpenAPS loop is set up in the wrong location on RasPi OS
[6 empty] | Incompatible OpenAPS and RasPi versions | OpenAPS developers shall specify the versions of RaspiOS which are support on each version of OpenAPS OpenAPS app shall do a compatibility check on whether it is compatible with the current RaspiOS version |  |  | OpenAPS files have been corrupted
[6 empty] | Human operator cancelled the RasPi set up | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
[6 empty] | Raspi OS crash | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
|  | UCA-STAO-3: Raspi OS does not start OpenAPS after there has been a completed update to OpenAPS | There has been an update to OpenAPS | Raspi OS does not start OpenAPS when it believes there has been a completed update to OpenAPS |  | Bug in OpenAPS code | Covered by Controller Constraint, CC-STAO-1: If there is a loop set up, when the Raspi OS is turned on it should start the OpenAPS loop | Raspi OS does not recognise that there has been a completed update to OpenAPS |  | OpenAPS files have been corrupted
[6 empty] | Incompatible OpenAPS and RasPi versions | OpenAPS developers shall specify the versions of RaspiOS which are support on each version of OpenAPS OpenAPS app shall do a compatibility check on whether it is compatible with the current RaspiOS version |  |  | OpenAPS does not communicate to RasPi OS that there has been a completed update and OpenAPS needs to run again
[6 empty] | Human operator cancelled the RasPi set up | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
[6 empty] | Raspi OS crash | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
|  | UCA-STAO-100: Raspi OS starts a different application and not OpenAPS | RasPi OS should start OpenAPS | Raspi OS starts a different application when it believes it should start OpenAPS |  | There is another file called OpenAPS | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active | Raspi OS believes that it should start a different application when it shold start OpenAPS |  | There is another instance of OpenAPS set as the default
[6 empty] | OpenAPS is set up in the wrong location | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active
|  | UCA-STAO-101: Raspi OS starts an older version of OpenAPS | The latest installed version of OpenAPS is version X (but Raspi OS starts version Y) | Raspi OS starts OpenAPS version Y when it believes that the latest version of OpenAPS installed is version X |  | Raspi OS bug | OpenAPS itself shall check to see if there is a newer version of OpenAPS installed | Raspi OS believes the latest OpenAPS version is version Y when it is version X |  | Configs show the wrong OpenAPS version and have not been updated
|  | UCA-STAO-102: Raspi OS starts OpenAPS twice | Raspi OS should only start OpenAPS once | Raspi OS starts OpenAPS twice when it believes that it should only start OpenAPS once |  | Raspi OS processer hangs as it is starting OpenAPS | OpenAPS shall check for multiple instances of it running | Raspi OS believes that it should start OpenAPS twice |  | -
UCA-STAO-103: Raspi OS starts OpenAPS loop when OpenAPS is not working properly
|  | UCA-STAO-1000: Raspi OS takes more than X seconds to start the OpenAPS loop | Raspi OS should have begun starting OpenAPS X seconds ago | Raspi OS takes more than X seconds to start the OpenAPS loop when it received the instruction to start OpenAPS X seconds ago |  | Kernel scheduler issues | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS | Raspi OS believes that it received the instruction to start OpenAPS more than X seconds ago |  | Timestamping problems, either with component clocks or Raspi OS clock
[6 empty] | Raspi OS is running slowly | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS
| Provide data from other devices to OpenAPS (RPD) | UCA-RPD-1: Raspi OS does not relay data from other devices to OpenAPS app when OpenAPS is running |  | Raspi OS does not relay data from other devices to OpenAPS app when it believes OpenAPS is running |  | RasPi OS does not receive data in correct format | If RasPi OS forwards the data in the incorrect format, OpenAPS shall check to ensure that the data is in the correct format. If RasPi OS does not forward the data, OpenAPS shall implement a timeout period to ensure it receives data in a timely manner. | Raspi OS does not recognise that OpenAPS is running |  | -
[6 empty] | RasPi OS does not finish receiving the data | If RasPi OS forwards the data in the incorrect format, OpenAPS shall check to ensure that the data is in the correct format. If RasPi OS does not forward the data, OpenAPS shall implement a timeout period to ensure it receives data in a timely manner. If either check fails, OpenAPS shall alarm the user.
[6 empty] | OpenAPS does not pull the data from Raspi OS | If OpenAPS fails to pull data from the RasPi OS, the user should be notified and the system should fallback to regular operation
|  | UCA-RPD-100: Raspi OS relays data from other devices to OpenAPS app inaccurately |  | Raspi OS relays data from other devices to OpenAPS app inaccurately when it initially received the correct data |  | RasPi OS received the data in the wrong format | (I need help with this one! It's not under OpenAPS's control) | Raspi OS received inaccurate data different from what other device sent |  | Data transmission from other devices back to the Raspi was interfered with
|  | UCA-RPD-101: Raspi OS relays data from other devices to OpenAPS app when other devices have not sent data to OpenAPS |  | Raspi OS relays data from other devices to OpenAPS app when it believes other devices have not sent data to OpenAPS |  | - |  | Raspi OS believes that other devices has sent data to OpenAPS when they have not |  | Data was sent to the RasPi OS from other components unintentionally
|  | UCA-RPD-1000: Raspi OS takes more than X seconds to relay data from other devices to OpenAPS app |  | Raspi OS takes more than X seconds to relay data from other devices to OpenAPS app when it received the data X seconds ago |  | Kernel scheduler issues | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS | RasPi OS believes that it received data from other device more than X seconds ago |  | Timestamping problems, either with component clocks or Raspi OS clock
[6 empty] | Raspi OS is running slowly | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS
| Stop OpenAPS (STOO) | UCA-STOO-1: Raspi OS does not stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes |  | Raspi OS does not stop OpenAPS when it believes OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes |  | OpenAPS does not inform the RasPi to stop OpenAPS when OpenAPS is not working properly | OpenAPS shall provide an error signal to the user indicating that something has gone wrong and attention is needed | Raspi OS is not informed that there have been suboptimal glucose leves for X minutes |  | Breakdown in communication between OpenAPS and the RasPi OS means RasPi OS did not receive the instruction to stop OpenAPS
[6 empty] | RasPi fails to stop OpenAPS completely and properly | OpenAPS shall provide an error signal to the user indicating that something has gone wrong and attention is needed
|  | UCA-STOO-100: Raspi OS stops OpenAPS when OpenAPS is running effectively and there has been no instruction from the human operator to stop OpenAPS |  | Raspi OS stops OpenAPS when it believes OpenAPS is running effectively and there has been no instruction from the human operator to stop OpenAPS |  | RasPi OS crash | The OpenAPS loop shall provide an "I'm alive!" signal whenever it is running so the user knows it is active | Raspi OS believes that there has been an instruction from the human operator to stop OpenAPS |  | Human operator inadvertently stops OpenAPS
|  | UCA-STOO-1000: Raspi OS takes more than X minutes to stop OpenAPS when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes |  | Raspi OS takes more than X minutes to stop OpenAPS when it believes OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes |  | Kernel scheduler issues | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS | RasPi OS does not recognise that it should have stopped OpenAPS already |  | Timestamping problems, either with component clocks or Raspi OS clock
[6 empty] | Raspi OS is running slowly | User shall have fresh install of Raspi OS on machine that meets minimum requirements with no other applications running except for OpenAPS
| Send command to Insulin Pump (SI) | UCA-SI-1: The RasPi OS does not send a command to the insulin pump when it has been sent a command by the OpenAPS app | RasPi OS has been sent a command by the OpenAPS app | RasPi OS does not send a command to the insulin pump when it believes it has been sent a command by the OpenAPS app |  | RasPi OS does not recognise the insulin pump | Device handshake from Raspi OS has to recognise insulin pump | RasPi OS does not believe that it has been sent a command by the OpenAPS app |  | Device had a automatic firmware update and the handshake config has to be updated on RasPi side | (1) Automatic firmware updates of the insulin pump have to be monitored by OpenAPS developers/community to make sure incompatibilities/need for user interaction are detected and communicated asap (2) OpenAPS user has to be aware of automatic device updates and the possibility of issues after an update (3) Alarm is needed if a handshake is expected but missing for too long (what is too long in this context?)
[6 empty] | RasPi OS does not have good network connection (airports / noisy areas) | Device must be kept from interference by many other signals and if the network connection falters the insulin pump should stop using OpenAPS and fallback to defaults
[6 empty] | RasPi OS does not believe that it is currently looping | RasPi OS must believe that it is looping when it is looping
[6 empty] | RasPi OS does not recognise the command | RasPi OS must be able to execute all the commands that OpenAPS app sends to RasPi OS
Algorithm does not believe that the command is safe
[6 empty] | RasPi OS crashes | Fallback
|  | UCA-SI-100: The RasPi OS sends a command to the insulin pump when it has not been sent a command by the OpenAPS app or a prompt from another component | RasPi OS has not been sent a command by the OpenAPS app or a prompt from insulin pump | The RasPi OS sends a command to the insulin pump when it believes it has not been sent a command by the OpenAPS app or a prompt from another component |  | Advesary hacks into RasPi and sends command to Insulin pump | RasPi OS and OpenAPS communication must be secure | RasPi OS believes that it has been sent a command by the OpenAPS app or a prompt from another component (when it hasn't) |  | There is a signal from somewhere else that RasPi OS mistakes as a signal from the OpenAPS app | (1) RasPi OS must be able to clearly identify the source of a command (2) Commands from OpenAPS to RasPi must be unique (3) RasPi OS and OpenAPS communication must be secure
[6 empty] | Someone has control of raspi by accident (maybe has controls ont their phone) to make it send a command | RasPi OS and OpenAPS communication must be secure
[6 empty] | Somebody's rig communicates with somebody else's insulin pump within the network | RasPi OS and OpenAPS communication must be secure
|  | UCA-SI-101: The RasPi OS sends the wrong command to the Insulin Pump | RasPi OS should send a different command to the Insulin pump | The RasPi OS sends the wrong command to the Insulin Pump when it believes it should send a different command to the insulin pump |  | Bug in RasPi OS command execution |  | RasPi OS believes that it should send the wrong command to the insulin pump |  | Intefering signals provide a different command to the RasPi OS to send the wrong command | (1) RasPi OS must be able to clearly identify the source of a command (2) Commands from OpenAPS to RasPi must be unique (3) RasPi OS and OpenAPS communication must be secure
[6 empty] | The function of a command has changed but this has not been updated in the RasPi OS |  |  |  | The wrong command is invoked to the RasPi | (1) OpenAPS commands must be compatible with RasPi OS (2) RasPi OS and OpenAPS communciation must be secure
|  | UCA-SI-102: The RasPi OS sends a command to the Insulin Pump when the Insulin Pump is not able to receive or action the command | Insulin pump is not able to receive or action the command | The RasPi OS sends a command to the Insulin Pump when it believes that the Insulin Pump is not able to receive or action the command |  | OpenAPS does not have the function to prevent the RasPi OS sending commands to the pump |  | RasPi OS believes that the Insulin Pump is not able to receive or action a command when it is |  | Bug in RasPi OS | (1) OpenAPS must recongnize when its communcation with RasPi OS is flawed (2) There needs to be an external entity to monitor behaviour of RasPi OS
The insulin pump has gone offline recently and OpenAPS sends the command to be queued for when it comes back online
RasPi OS recognized that the insulin pump went offline too late and still sends a command (bug)
|  | UCA-SI-1000: The RasPi OS sends a command to the Insulin pump more than X seconds after the command has been invoked by OpenAPS | OpenAPS invoke an alias command at least X seconds ago | The RasPi OS sends a command to the Insulin pump more than X seconds after the command has been invoked by OpenAPS when it believes that OpenAPS invoked its command at least X seconds ago |  | Kernel scheduler issue | Use independent or external monitoring software that can detect unexpected or unusual operation in the kernel When there are scheduling issues, pump must fall back to regular operation | RasPi OS incorreclty believes that OpenAPS invoked it's command to the insulin pump less than X seconds ago |  | Timestamping issue | Timestamp to include error checking, redundancies, or counter to reduce the likelihood of faulty timestamp
[6 empty] | Command or function in algorithm takes too long or sends too much data | Use independent or external monitoring software that can detect unexpected or unusual operation in the kernel When there are scheduling issues, pump must fall back to regular operation |  |  | Different times on OpenAPS rig and other components
| Send command to NightScout (SN) | UCA-SN-1: The RasPi OS does not send a command to NightScout when it has been sent a command by the OpenAPS app | RasPi OS has been sent a commnad by the OpenAPS app | The RasPi OS does not send a command to NightScout when it believes that it has been sent a command by the OpenAPS app |  | RasPi OS does not recognise NightScout | RasPi OS should be able to recognise all devices that OpenAPS is expecting | RasPi OS incorrectly believes that it has not been sent a command by the OpenAPS app |  | Device had a automatic firmware update and the handshake config has to be updated on RasPi side | Include version checking in the apps that will warn user to update the Raspbian OS version if this occurs
[6 empty] | RasPi OS does not have good network connection (airports / noisy areas) | Either: If a command from OpenAPS is not received by an external devices (eg for network reasons), if not safety-critical place command in queue to resend later If a command from OpenAPS is not received by an external devices (eg for network reasons), if safety-critical warn the user and fallback to regular operation
[6 empty] | RasPi OS does not believe that it is currently looping | Implement flag structure in loop that will only get set when it is looping and warn user/shut down if not looping (may need external monitoring device)
[6 empty] | RasPi OS does not recognise the command | RasPi OS should recognise all commands that it has been issued by the OpenAPS app
[6 empty] | Algorithm does not believe that the command is safe | If the algorithm determines that a command is unsafe, the system should fallback to regular insulin pump operation and the user should be notified
[6 empty] | RasPi OS crashes | Implement external monitoring device that can detect crash and reset Raspberry Pi
|  | UCA-SN-100: The RasPi OS sends a command to NightScout when it has not been sent a command by the OpenAPS app or a prompt from the insulin pump | RasPi OS has not been sent a command by the OpenAPS app or a prompt from another component | The RasPi OS sends a command to NightScout when it believes it has not been sent a command by the OpenAPS app or a prompt from the insulin pump |  | Adversary hacks into RasPi and sends command | Industry-standard security techniques should be used to prevent hacking Monitoring software that can detect unexpected or dangerous commands should be implemented | RasPi OS incorrectly believes that it has been sent a command by the OpenAPS app |  | There is a signal from somewhere else that RasPi OS mistakes as a signal from the OpenAPS app | Include unique identifiers (address, application name etc.) to more clearly indicate if a signal is from the OpenAPS app
[6 empty] | Someone has control of raspi by accident (maybe has controls ont their phone) to make it send a command | Implement unique identifier or two-step verification prompt when there's a new input device that prevents accidental sending of commands from an unexpected source
[6 empty] | Rig communicates with somebody else's insulin pump within the network | Covered
|  | UCA-SN-101: The RasPi OS sends the wrong command to NightScout | RasPi OS should send a different command to the other component than what it did | The RasPi OS sends the wrong command to NightScout when it believes it should send a different command |  | Bug in RasPi OS command execution | If Raspi OS is malfunctioning and not sending the correct commands, the user should be notified to reinstall the Raspi OS and the piu | RasPi OS believes that it should send the wrong command to NightScout |  | Intefering signals provide a different command to the RasPi OS to send the wrong command | Include error-checking or redundancies in all commands that will reduce the likelihood of interference causing a misinterpretation of the command
[6 empty] | The function of a command has changed but this has not been updated in the RasPi OS | Compatible versions of Raspbian OS with OpenAPS shold make sure all the latest openaps packages are installed |  |  | The wrong command is invoked to the RasPi
|  | UCA-SN-102: The RasPi OS sends a command to NightScout when the other device is not able to receive or action the command | Other component is not able to receive or action the command | The RasPi OS sends a command to NightScout when it believes the other device is not able to receive or action the command |  | The component has gone offline recently and OpenAPS sends the command to be queued for when it comes back online | OpenAPS should send commands to be queued, wait for acknowledgement from NightScout; otherwise keep in queue and attempt to resend again | RasPi OS incorrectly believes that the other device is ablt to receive or action the command |  | Bug in RasPi OS | If there is a bug in Raspberry Pi OS, the user should be warned, the system should fallback to regular operation and that version of Raspberry Pi OS should be marked as incompatible in the OpenAPS documentation
[6 empty] | RasPi OS recognized that the component went offline too late and still sends a command (bug) | Wait for acknowledge from NightScout; otherwise keep in queue and attempt to resend again |  |  | OpenAPS does not have the ability to detect that NightScout is not able to receive or action the command from OpenAPS | OpenAPS should be able to detect whether the other component is able to receive and action command
|  | UCA-SN-1000: The RasPi OS sends a command to NightScout more than X seconds after the command has been invoked by OpenAPS | OpenAPS invoke an alias command at least X seconds ago | The RasPi OS sends a command to NightScout more than X seconds after the command has been invoked by OpenAPS when it believes that OpenAPS invoked its command at least X seconds ago |  | Kernel scheduler issue | Use independent or external monitoring software that can detect unexpected or unusual operation in the kernel When there are scheduling issues, pump must fall back to regular operation | RasPi OS incorreclty believes that OpenAPS invoked it's command to the other device less than X seconds ago |  | Timestamping issue | Timestamp to include error checking, redundancies, or counter to reduce the likelihood of faulty timestamp
[6 empty] | Command or function in algorithm takes too long or sends too much data | Use independent or external monitoring software that can detect unexpected or unusual operation in the kernel When there are scheduling issues, pump must fall back to regular operation |  |  | Different times on OpenAPS rig and other components | Standardize timesource or use timesource separate from the main system clocks
OpenAPS app | Issue building block command to RasPi OS for communicating with another device (ICP) | UCA-ICP-1: OpenAPS app does not invoke the corresponding building block commands to the RasPi OS when the algorithm invokes an alias command to OpenAPS app | Algorithm has invoked an alias command to OpenAPS app | OpenAPS app does not invoke the corresponding building block commands to the RasPi OS when it believes the algorithm has invoked an alias command to OpenAPS app |  | Process(or) is hung | When RasPi processor is hung, insulin pump must fall back to regular operation | Algorithm has invoked an alias command to OpenAPS app but OpenAPS app does not realise this |  | Mismatch between OpenAPS app latest version and version algorithm is using
[6 empty] | Memory or other hardware failure | When RasPi hardware fails, insulin pump must fall back to regular operation |  |  | Algorithm did not pull full list of OpenAPS app commands correctly
[6 empty] | Bug in algorithm command invoking communication |  |  |  | Algorithm did not trigger OpenAPS app command correctly
[6 empty] | OpenAPS app repository location might be moved / different in relation to algorithm repository location | The algorithm must successfully import commands from the OpenAPS app
[6 empty] | OpenAPS cannot locate receiving device on the network | -
|  | UCA-ICP-100: OpenAPS app invokes corresponding commands to the RasPi OS when the Algorithm has not invoked an alias command to the OpenAPS app | Algorithm has not invoked an alias command to the OpenAPS app | OpenAPS app invokes corresponding commands to the RasPi OS when it believes the Algorithm has not invoked the alias command to the OpenAPS app |  | - |  | Algorithm has not invoked alias command to OpenAPS app but OpenAPS app believes that it has |  | ?
|  | UCA-ICP-101: OpenAPS app invokes commands to be used to the RasPi OS when the Algorithm has invoked an alias command to OpenAPS app that do not correspond to the command that the OpenAPS app issued to the RasPi OS | Algorithm has invoked an alias command to OpenAPS app that does not correspond to a command that the OpenAPS app issued by the RasPi OS | OpenAPS app invokes commands to be used to the RasPi OS when the Algorithm has invoked an alias command to OpenAPS app that do not correspond to the command that the OpenAPS app issued to the RasPi OS |  | - |  | OpenAPS app believes that the Algorithm has invoked a different command then it has |  | ?
|  | UCA-ICP-1000: OpenAPS invokes corresponding commands to the RasPi OS more than X seconds after the alias command has been used by the algorithm | Alias command has been used by the algorithm at least X seconds ago | OpenAPS app invokes corresponding commands to the RasPi OS more than X seconds after the alias command has been used by the algorithm when it believes the alias command has been used by the algorithm at least X seconds ago |  | Kernel scheduler issue | Covered above | OpenAPS app believes that the command was used less than X seconds ago |  | Timestamping issue | Covered above
[6 empty] | Function in algorithm takes too long or sends too much data | Watchdog must be in place to monitor individual functions in the algorithm Function in the algorithm must take less than X seconds to be applied to the RasPi OS |  |  | Different times on OpenAPS rig and other components | Covered above
| Set up algorithm and loop (OSA) | UCA-OSA-1: OpenAPS app does not update the configurations (or prompt the human operator to do so) when there is a change in the OpenAPS loop setup that should affect the configurations | There has been a change in the OpenAPS loop setup that should affect the configurations | OpenAPS app does not update the OpenAPS configurations (or prompt the human operator to do so) when it believes there is a change in the OpenAPS loop setup that should affect the configurations | User Interface believes there has been a change in the OpenAPS loop setup that should affect the configurations | There is no information in the code/configuration that allows for automated compatibility checks | - | There has been a change in the OpenAPS loop setup that should affect the configurations but OpenAPS does not recognise this |  | If there is information in the code/configuration that allows compatibility checks, the version information is not updated, so that the compatiblity algorithm thinks everything is still on its old configuration | -
[6 empty] | If there is information in the code/configuration that allows compatibility checks, the compatibility checks are not implemented or errorneously implemented | Any compatibility check relating to OpenAPS configurations must accurately report when the configurations are currently not accurate |  |  | There have been additional configurations added in the latest OpenAPS version so the version that is being used needs to be updated | When the user updates their version of OpenAPS, they should be prompted to set any new config settings The OpenAPS documentation should list any new config settings that need to be set for each new version When the user updates the version of OpenAPS, they should run the setup script again
[6 empty] | The interface detects the configuration mismatch, but does not show this on the UI | If OpenAPS detects that one of the config values are not optimally set, it must indicate this on the UI
[6 empty] | OpenAPS is not set up to prompt the human to update configs when it detects a change that should be reflected in configs or detects inaccurate configs | If OpenAPS detects that one of the config values are not optimally set, it must indicate this on the UI
|  | UCA-OSA-2: OpenAPS app does not successfully set up the algorithm when the human operator runs the script to set up the algorithm | Human operator has run the script to set up the algorithm | OpenAPS app does not successfully set up the algorithm when it believes that the human operator runs the script to set up the algorithm |  | Dependencies on the RaspiOS are not / cannot be the correct version to install the algorithm correctly | For each RasPI OS dependency, the version(s) that are compatible each supported version of OpenAPS must be documented | Human operator has run the script to set up the algorithm butOpenAPS statuses do not indicate this |  | Installation script aborts or crashes partway through | Human operator must have access to the command line so they can see whether the installation script suceeded Algorithm installation script should be tested on all versions of OpenAPS that expected to be supported and the results should be feed into the documentation supported versions
[6 empty] | OpenAPS does not support RaspiOS software or hardware version or something is interfering from the host side | For each RasPI OS dependency, the version(s) that are compatible each supported version of OpenAPS must be documented If there are accessories of the hardware that intefere with OpenAPS or the drivers that OpenAPS uses, this should be documented.
[6 empty] | OpenAPS has packaging problems / broken packages which are causing problems installing dependencies | OpenAPS must not have broken dependencies and packages
[6 empty] | Bug in the script to run the algorithm | OpenAPS install script must run successfully for all valid setups of OpenAPS
|  | UCA-OSA-100: OpenAPS app selects a different option to that which the human operator entered or intended when configuring OpenAPS | Human operator entered or intended a different configuration option than is selected | OpenAPS selects a different option to that it believes the human operator entered or intended when configuring OpenAPS | User Interface believes the human operator entered or intended a different option | OpenAPS bug or malfunction | OpenAPS must differentiate between user-enacted configurations and system-enacted configurations and must prioritze the user configurations over system configurations | Human operator entered or intended a different configuration option than is selected but the OpenAPS information does not indicate this |  | Issue with recognizing the operator input
selection of a different memory section than the one the information entered by the operator is stored at (memory address issue)
[6 empty] | Issue with passing the entered information in OpenAPS | OpenAPS must confirm the entered information (as stored/used) with the user
[6 empty] | Information entered by the human is not valid and OpenAPS has defaulted to a valid setting | OpenAPS must communicate to the user that the entered information is not valid; OpenAPS must communicate clearly what options are being used for operations.
|  | UCA-OSA-101: OpenAPS sets up an unsafe algorithm | Algorithm is not optimal for OpenAPS | OpenAPS sets up an algorithm that it recognises is unsafe | User Interface believes algorithm is non-optimal | OpenAPS bug or malfunction | OpenAPS must set up the algorithm that the human has entered in the setup script | Algorithm is not optimal for OpenAPS but the OpenAPS app information does not indicate this |  | Algorithm has safety flaws that noone knows about | OpenAPS should test and track an algorithms performance with a wide range of different config settings OpenAPS must have a forum where people can inform developers of algorithm malfunctions
[6 empty] | Human confirms that the algorithm is safe when it has not been tested properly | Human must make sure that algorithm has been tested properly before using it OpenAPS must specify in the documentation how to make sure an algorithm is safe OpenAPS should provide a service to test the safety of algorithms OpenAPS should have a facility to allow the human to test an algorithm, while running a different algorithm, and compare dosing recommendations |  |  | OpenAPS misidentifies the algorithm as a recommended algorithm | OpenAPS must set up the algorithm that the human has entered in the setup script
[6 empty] | OpenAPS does not stop the human from setting up unsafe algorithms | Covered above |  |  | Algorithm is known to be unsafe but there is no feature on OpenAPS to blacklist this | Covered
[6 empty] | OpenAPS has a mode which allows any algorithm to be used | OpenAPS should have a facility to allow the human to test an algorithm, while running a different algorithm, and compare dosing recommendations
|  | UCA-OSA-1000: OpenAPS app updates the algorithm or configurations too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings | OpenAPS has already started to operate with the current version | OpenAPS app updates too late, after OpenAPS has already started to operate, and is instead using previous algorithm settings, when it believes that OpenAPS has already started with previous algorithm settings | User Interface believes that it's too late for the algorithm to encorporate the update | Raspi OS bug or malfunction | When new settings have been set, OpenAPS must update based on these settings before the next loop | OpenAPS has already started to operate with the current version but the OpenAPS information does not indicate this |  | OpenAPS does not ensure that the configs are updated before running the algorithm | OpenAPS should have a fixed spot to check for and implement updates, either new settings or software updates, in every (~5 minute) OpenAPS loop
[6 empty] | Some other action is blocking the update routine | OpenAPS should have a fixed spot to check for and implement updates in every (~5 minute) OpenAPS loop
[6 empty] | Scheduling/Interrupt issue | Covered above
| Provide data from other devices to algorithm (OPD) | UCA-OPD-1: OpenAPS app does not relay relevant data from other devices to the algorithm |  | OpenAPS does not relay data from other devices to algorithm when it believes the data is relevant |  | OpenAPS does not receive data in correct format |  | OpenAPS does not believe the data is relevant to the algorithm |  | Bug or lack of complexity in the algorithm | (1) Data formats (APIs ?) that OpenAPS works with must be specified (2) If data that has an unexpected data format/data content is received, OpenAPS must be able to detect it (3) There must be a defined reaction of OpenAPS to unexpected data (to alarm the user or the other device/s) (4) Data faults that cannot be detected by OpenAPS must be avoided or mitigated by the other devices
[6 empty] | OpenAPS does not finish receiving the data |  |  |  | OpenAPS does not recognise all the inputs that the algorithm uses | There must be verification measures to ensure that all used interfaces are defined and implemented
[6 empty] | OpenAPS does not pull the data from Raspi OS |  |  |  |  | There must be a control flow monitoring to ensure all OpenAPS routines are run as planned.
|  | UCA-OPD-100: OpenAPS app relays data from other devices to algorithm inaccurately |  | OpenAPS app relays data from other devices to OpenAPS app inaccurately when it initially received the correct data |  | OpenAPS app received the data in the wrong format |  | OpenAPS received inaccurate data different from what RaspiOS sent |  | Bug in Raspi data processing | Details of data processing must be specified to ensure the processing algorithms are suitable for the data formats, sizes and timing Correct specification and implementation of data processing must be verified
|  | UCA-OPD-101: OpenAPS app relays data from other devices to algorithm when other devices have not sent data to OpenAPS |  | OpenAPS app relays data from other devices to algorithm when it believes other devices have not sent data to OpenAPS |  | - |  | OpenAPS believes that other devices has sent data to algorithm when they have not |  | Data was sent to the OpenAPS from Raspi OS unintentionally
|  | UCA-OPD-1000: OpenAPS app takes more than X seconds to relay data from other devices to the algorithm |  | OpenAPS app takes more than X seconds to relay data from other devices to algorithm when it received the data X seconds ago |  | Kernel scheduler issues |  | OpenAPS believes that it received data from Raspi OS more than X seconds ago |  | Timestamping problems, either with component clocks or Raspi OS clock | (1) OpenAPS must be able to detect timed-out data (2) OpenAPS must have defined behaviour when timed-out data is detected
[6 empty] | Raspi OS is running slowly |  |  |  |  | OpenAPS must be able to detect gaps in the detected data (for detecting that data packages sent from RaspbiOS are lost)
[6 empty] | OpenAPS script has been updated and now takes too long |  |  |  |  | OpenAPS must be able to detect gaps in the detected data (for detecting that data packages sent from RaspbiOS are lost)
| Stop Algorithm (STOA) | UCA-STOA-1: OpenAPS app does not stop algorithm and fall back to regular pump operation when OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes |  | OpenAPS does not stop the algorithm when it believes OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for X minutes |  | Bug in OpenAPS means that OpenAPS fails to stop the algorithm completely and properly |  | OpenAPS is not informed that there have been suboptimal glucose levels for X minutes |  | CGM is inaccurate
[6 empty] | OpenAPS thinks that it can fix the problem itself by altering dosing |  |  |  | There is a problem with communication to OpenAPS
OpenAPS thinks that the problem is resolved and prematurely starts to give dosing recommendations again
|  | UCA-STOA-100: OpenAPS app stops algorithm when OpenAPS is giving good dosing recommendations to the insulin pump and the system is running correctly and there has been no instruction from the human operator to stop OpenAPS |  | OpenAPS stops the algorithm when it believes OpenAPS is running effectively and there has been no instruction from the human operator to stop OpenAPS |  | OpenAPS crash |  | OpenAPS believes that there has been an instruction from the human operator to stop the algorithm
|  | UCA-STOA-1000: OpenAPS app takes more than X minutes stop algorithm when the algorithm is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes |  | OpenAPS app takes more than X minutes to stop algorithm when it believes OpenAPS is giving dosing recommendations which are causing suboptimal glucose levels for Y minutes |  | Kernel scheduler issues |  | OpenAPS app does not recognise that it should have stopped the algorithm already |  | Timestamping / clock issues
Breakdown in communication between OpenAPS and the RasPi OS means RasPi OS did not receive the instruction to stop OpenAPS
Human operator inadvertently stops OpenAPS
Timestamping problems, either with component clocks or Raspi OS clock
Algorithm | Get aliases and commands to use from OpenAPS app | UCA-GCO-1: Algorithm does not pull an OpenAPS app command, that it uses, from the OpenAPS app | OpenAPS app command is useful for the algorithm | Algorithm does not pull an OpenAPS app command, that it believes it uses, from the OpenAPS app | Algorithm believes that the command is useful | A command has been named differently in the algorithm and in the OpenAPS app | The names of OpenAPS app commands that the algorithm uses should be all present in the version of the OpenAPS app that is being used | OpenAPS app command is useful for the algorithm but the algorithm information does not indicate this |  | An OpenAPS app command has been removed or renamed | When an updated version of OpenAPS app is being used, the app commands in the OpenAPS algorithm should be refactored as necessary There should be a notification to the user that the OpenAPS app commands have been updated
[6 empty] | The location of the OpenAPS app or of the algorithm has been moved | The OpenAPS app commands must be imported successfully by the algorithm (this is a Controller constraint)
[6 empty] | OpenAPS rig stops working when the algorithm is trying to pull a command from the OpenAPS app | OpenAPS should feedback error information when the RasPi rig stops working in the middle of a command
|  | UCA-GCO-2: Algorithm does not pull useful changes and additions that have been made to OpenAPS app commands | There have been useful additions made to the OpenAPS app commands | Algorithm does not pull useful changes and additions that it believes have been made to OpenAPS app commands | Algorithm believes useful changes and additions have been made to OpenAPS app commands | A command has been named differently in the algorithm and in the OpenAPS app | The names of OpenAPS app commands that the algorithm uses should be all present in the version of the OpenAPS app that is being used | There have been useful additions made to the OpenAPS app commands but the algorithm information does not indicate this |  | There is no information to the algorithm or the developer when the app has been updated | There should be a notification to the user that the OpenAPS app commands have been updated
[6 empty] | The location of the OpenAPS app or of the algorithm has been moved | The OpenAPS app commands must be imported successfully by the algorithm (this is a Controller constraint)
[6 empty] | OpenAPS rig stops working when the algorithm is trying to pull a command from the OpenAPS app | OpenAPS should feedback error information when the RasPi rig stops working in the middle of a command
|  | UCA-GCO-100: Algorithm pulls aliases from the OpenAPS app with a different name to what that command is called in the algorithm | OpenAPS app command is called a (different) name | Algorithm pulls aliases from the OpenAPS app with a different name to what it believes that command is called in the algorithm | The algorithm believes that the command has a different name | There is a typo in the algorithm and the OpenAPS tool name is spelt wrong | Covered | OpenAPS app command is called a (different) name but the algorithm information does not indicate this |  | The name of the OpenAPS tool has been changed but the algorithm has not been informed of this | Covered
[10 empty] | There is a typo in the algorithm and the OpenAPS tool name is spelt wrong | Covered
|  | UCA-GCO-101: Algorithm pulls the same command from the OpenAPS app more times than necessary | It should only be necessary for the algorithm to pull the command from the OpenAPS app once | Algorithm pulls the same command from the OpenAPS app more times than necessary | Algorithm believes it is only necessary to pull the command from the app once / (x times) | - |  | It should only be necessary for the algorithm to pull the command from the OpenAPS app once but the algorithm information does not indicate this |  | -
|  | UCA-GCO-1000: Algorithm pulls aliases and commands from the OpenAPS app after they are first used in the algorithm | Aliases/commands from the OpenAPS app have already been used in the algorithm | Algorithm pulls aliases and commands from the OpenAPS app when it believes they have been used in the algorithm | Algorithm believes that the alias/command has already been used in the algorithm | Algorithm starts before aliases and commands have finished being pulled | All aliases and commands should be pulled before the algorithm begins | Aliases/commands from the OpenAPS app have already been used in the algorithm but the algorithm information does not indicate this |  | The algorithm has been updated to use the latest commands but the latest commands have not been pulled in | Covered
| Issue command to OpenAPS app to communicate with other devices (ICO) | UCA-ICO-1: Algorithm does not invoke a command to the OpenAPS app when that command is used in the algorithm and it's an OpenAPS app command | OpenAPS app command is used in the algorithm | Algorithm does not invoke a command to the OpenAPS app when it believes that command is used in the algorithm and it's an OpenAPS app command |  | Process(or) is hung | When RasPi processor is hung, insulin pump must fall back to regular operation | OpenAPS app command is used in the algorithm but the algorithm information does not indicate this |  | Algorithm interferes with OpenAPS app command names | Algorithm must not interfere with OpenAPS app command names
[6 empty] | Memory or other hardware failure | When RasPi hardware fails, insulin pump must fall back to regular operation |  |  | Mismatch between OpenAPS app latest version and version algorithm is using | OpenAPS must make version check of OpenAPS app being used against latest version available OR human user must be notified by email when there is a new OpenAPS app version
[6 empty] | Bug in algorithm command when invoking communication | Components must be connected to each other When communication is failing, insulin pump must fall back to regular operation |  |  | Algorithm did not pull full list of OpenAPS app commands correctly | Covered
[6 empty] | OpenAPS app repository location might be moved / different in relation to algorithm repository location | The algorithm must successfully import commands from the OpenAPS app
|  | UCA-ICO-100: Algorithm invokes a command to the OpenAPS app which is not an OpenAPS app command | Command is not an OpenAPS app command | Algorithm invokes a command to the OpenAPS app which it believes is not an OpenAPS app command |  | Algorithm has a bug which makes it apply a non OpenAPS app command | Algorithm must only use OpenAPS app commands to communicate with other devices | Command is not an OpenAPS app command but the algorithm information does not indicate this |  | Function in the algorithm shares a name with a function in the OpenAPS app | Function in the algorithm must not share a name with a function in the OpenAPS app
[10 empty] | Algorithm interferes with OpenAPS app command names | Covered
[10 empty] | Mismatch between OpenAPS app latest version and version algorithm is using | Covered
[10 empty] | Algorithm did not pull full list of OpenAPS app commands correctly | Covered
|  | UCA-ICO-101: Algorithm invokes a command to the OpenAPS app which is not used in the algorithm at that time | Command is not used in the algorithm at the time | Algorithm invokes a command to the OpenAPS app that it believes is not used in the algorithm |  | We don't think this can happen | - | Command is not used in the algorithm at the time but the algorithm information does not indicate this |  | We don't think this can happen | -
|  | UCA-ICO-102: Algorithm invokes the wrong command or a deprecated command to the OpenAPS app | Command is incorrect/deprecated | Algorithm invokes a command to the OpenAPS app that the algorithm believes is a wrong command or a deprecated command |  | Algorithm has typo or bug which makes it apply the wrong command | Covered | Command is incorrect / deprecated but the algorithm information does not indicate this |  | Update to app that changes or deprecates a command | Covered
[10 empty] | List of app commands is not correct | Covered
[10 empty] | No notifications / communication between developers when updating APIs | There must be communication about updates between developers of different but related repositories
[10 empty] | APIs between algorithm and app have changed and they are not synced | Covered
|  | UCA-ICO-103: Algorithm uses an OpenAPS app command in a wrong way | Command should be used differently | Algorithm uses an OpenAPS app command in one way when it believes that the OpenAPS command should be used differently |  | Bug in the algorithm |  | Command should be used differently but the algorithm information does not indicate this |  | Lack of clear documentation around invocation of command by OpenAPS app developer | OpenAPS app commands must be clearly documented
[10 empty] | Commands named too similarly | -
[10 empty] | Command has changed it's behaviours between versions | Developers should be informed about which commands have changed behaviours in different versions
|  | UCA-ICO-1000: Algorithm invokes a command to the OpenAPS app more than X seconds after the command comes up in the algorithm | Command has come up in the algorithm at least X seconds ago | Algorithm invokes a command to the OpenAPS app more than X seconds after the command comes up in the algorithm, when it believes that the command has come up in the algorithm at least X seconds ago |  | Kernel scheduler issue |  | Command has come up in the algorithm at least X seconds ago but the algorithm information does not indicate this |  | Timestamping issue
[6 empty] | Command or function in algorithm takes too long or sends too much data |  |  |  | Different times on OpenAPS rig and other components | OpenAPS should warn user when other components have different timestamps and also refactor the incoming timestamps of other components in this case
Performance issue
Processor is overloaded

## Tracing commands control actio
Scenario number | Command | Context
1 | openaps init <name> (Creates an instance of OpenAPS) | Specify and set up algorithm and loop
2 | tail -F /var/log/openps/pump-loop.log (shows the log of the current running instance of OpenAPS, output should be long if the OpenAPS instance is working properly) | Human Start openaps and feedback line, Report operating data
3 | openaps use pump read_clock (read pump date/time) | OpenAPS System request operating data from insulin pump
4 | openaps use pump model (read pump model number) | OpenAPS System request operating data from insulin pump
5 | openaps use pump read_basal_profile_A (read pump basal profile A) | OpenAPS System request operating data from insulin pump
6 | openaps use pump suspend_pump (suspend pumping) | RasPi OS send instruction to Insulin pump

## Questions, Variables, Assumptio
Questions | Formatting | Assumptions | Variables to be assigned values | Questions
What is the middleware running in Pi | If a requirement or scenario has a variable to be assigned a value, add it to the list on the right | The other FDA Approved devices in the system, such as Medtronic Insulin pumps and Dexcom CGMs, are safe and reliable | Maximum time that a non optimal dosing recommendation should stand without being updated | What is the middleware running in Pi
| If you see a problem with a loss scenario on the Requirements special page, change it on the requirements special page rather than the loss scenario page | The BUS which is being used for communication is secure and trustworthy, not liable to data loss or corruption | Maximum number of minutes without getting new data from the insulin pump
Frequency of infomation from insulin pump? (5 mins) | From now if you want to add a new loss scenario, first search for the keywords on the Loss Scenario Analysis page | Human is using oref0 algorithm | Maximum number of minutes without getting data from the NightScout | Frequency of infomation from insulin pump? (5 mins)
Does OpenAPS know there's new information, or does it pull periodically? | Loss Scenarios should be written in present tense, avoid starting the loss scenarios with "The" | Currently we are considering the case of NightScout being used, Autotune not being used | Minimum number of seconds gap between dosing recommendation issued and last insulin pump data pull | Does OpenAPS know there's new information, or does it pull periodically?
What services does OpenAPS / Pi use? | OpenAPS Algorithm Loss Scenarios should be one of : Flawed OpenAPS decision making algorithm (in this case specify UCA) Flawed OpenAPS data analysis Flawed OpenAPS dosing calculation algorithm |  | Maximum number of data requests from pump a minute
Maximum number of data requests from CGM a minute
What networking does OpenAPS use | There may be more than one requirement for one Loss Scenario |  | Maximum number of weeks running old CGM software
Maximum time between question asked on forum and being answered
| Have Controller Constraint UCAs there. If there is a Controller Constraint UCA, then you don't need to make a Requirement for it from a Scenario |  | CGM clock shold match RasPi time and real clock time to the nearest X seconds
| Please feel free to take a Loss Scenario from the Requirements page and 'Find' it on the Loss Scenarios page, that way you'll be able to count all the ways that the Loss Scenarios were grooving |  | Timestamps on data should be accurate (to whichever clock) to X seconds
| (Include a Loss Scenario Count) |  | OpenAPS must get valid data from the NightScout when it has not received data which is from the last X minutes
OpenAPS must get data from NightScout at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
OpenAPS must not get data from the CGM more than X times a minute, if NightScout is not available (consequence: could diminish battery or wear out the CGM)
OpenAPS must get data from the insulin pump when it has not received data which is from the last X minutes
| Different UCAs have the same Loss Scenario |  | OpenAPS must get data from the insulin pump at least X seconds before the next dosing recommendation, when X seconds are needed to incorporate the data into the next dosing recommendation
| Different Loss Scenarios having the same requirement |  | OpenAPS must not request data from the insulin pump more than X times a minute
| Loss Scenario has multiple requirements |  | OpenAPS must set the temporary basal dose when the current insulin dosage is not optimal and it has been over X minutes since the last temporary basal recommendation
Human actor must provide meal data to the insulin pump within X minutes from the actual mealtime
OpenAPS should make a notification when the insulin pump settings do not match the recommended configuration in the OpenAPS documentation with X frequency
Make sure human knows the time for the insulin to react in their body and sets the advanced time (by X minutes) accordingly
(Note, Is it possible for device to be 'conversive' and react to changes in glucose levels with questions about recent human activity, also note fast acting insulin is preferred)
OpenAPS should warn the human actor if human has not input useful daily information in X hours
OpenAPS must check every X minutes for all the latest information
OpenAPS must not use information older than X minutes for dosing recommendation
OpenAPS should prompt the human for data if glucose levels are X mg lower than expected
The time between OpenAPS being sent the data to make the dosing calculation and OpenAPS issuing and insulin dose must be < X seconds
If possible, OpenAPS must monitor insulin pump battery

## Common Scenarios
Algorithm delay
Kernel scheduler issue
OpenAPS algorithm takes too long
Processor hanging / overloaded
Device overheating
General system problems or host crash
Low voltage / poor power supply
Raspi battery runs out
OpenAPS clock set to the wrong time (changing timezone, clocks changing, clock error over time)
The algorithm has been changed and something now takes longer, but the timing of when things happen has not been adjusted to accomodate for this
General poor task scheduling by OpenAPS
Another component sends data with the wrong timestamp
Human operator activity on device
Host lack of memory space causing poor performance, this also may stop logs from being written
Dependency issues / system update
Wifi latency causes communication issues which leads to delays
Question: Does OpenAPS have a detection mechanism to flag whether algorithm issues are delays or bugs, and push alert the user when there is a serious issue? There should ideally be a push for sure. (Add to requirements)
Question: How does OpenAPS ensure time changes don't compromise the dosage, if doses are scheduled for a time?
Question: How are OpenAPS and other components deriving time? Raspi uses NTP server for time. User pump / components may have manual clock setting. How does time in other components affect the OpenAPS algorithm?
Question: How does OpenAPS ensure that it still has enough space to run, and save logs? Does it manage disk space? Does it roll over / delete logs?
Question: How does OpenAPS handle low battery? How long does the CHO stress test last? How do they perform this test / what do they test?

## Unused analysis
MISSING!!!!!!!!!!!!!!!!!!!!!!!!
Provide data or dosing instructions directly to OpenAPS rig, via a phone app ?
OBSELETE!!!!!!!!!!!!!!!!!!!!
|  |  | UCA-SB-101 : OpenAPS algorithm recommends a too low bolus dosing recommendation based on current glucose levels, active insulin and human insulin pump input |  | UCA-SB-1001 : OpenAPS provides a bolus recommendation too soon after the last bolus dose from the insulin pump
UCA-SB-102: OpenAPS makes a bolus recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
UCA-SB-103 : OpenAPS makes a bolus dosing recommendation which overrides the human’s safe requested level
UCA-SB-104 : OpenAPS provides a bolus dosing rate when the insulin pump settings are not configurated correctly (*1)
UCA-SB-105: OpenAPS makes a bolus recommendation when it is running an algorithm that should only set temporary basal rates
UCA-SB-106 : OpenAPS makes bolus recommendations too frequently (repetitive stuck)
UCA-STB-107 : OpenAPS makes a bolus recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
Set Bolus (SB) | UCA-SB-1: OpenAPS does not set a supermicrobolus for a meal when it is running the oref1 algorithm | OpenAPS is running the oref1 algorithm | OpenAPS believes it is running a non-bolus algorithm
| UCA-SB-100 : OpenAPS algorithm recommends a too high bolus dosing recommendation based on current glucose levels, active insulin and human insulin pump input |  |  |  | ...
UCA-SB-101 : OpenAPS algorithm recommends a too low bolus dosing recommendation based on current glucose levels, active insulin and human insulin pump input
UCA-SB-102: OpenAPS makes a bolus recommendation when there is inadequate data available to it to justify the dosing recommendation (eg not enough data, data is too conflicting)
UCA-SB-103 : OpenAPS makes a bolus dosing recommendation which overrides the human’s safe requested level
UCA-SB-104 : OpenAPS provides a bolus dosing rate when the insulin pump settings are not configurated correctly (*1)
UCA-SB-105: OpenAPS makes a bolus recommendation when it is running an algorithm that should only set temporary basal rates
| UCA-SB-106 : OpenAPS makes bolus recommendations too frequently |  |  | OpenAPS is using a flawed algorithm which is causing too frequent bolus recommendations
UCA-SB-107 : OpenAPS makes a bolus recommendation when the insulin pump is not in a state to give a dose (no battery, no insulin, infusion set needs changing)
| UCA-SB-1000 : OpenAPS provides a bolus recommendation too late after a meal time | A meal was more than X minutes ago | OpenAPS believes that the meal was less than X minutes ago | OpenAPS processing delays | OpenAPS is configured with different timestamps to the insulin pump and believes that it has dosed the insulin within X minutes of the meal
|  |  | OpenAPS believes that it has provided a dosing recommendation | Linux kernel scheduling has caused OpenAPS to dose bolus later than X minutes after the meal
| UCA-SB-1001 : OpenAPS provides a bolus recommendation too soon after the last bolus dose from the insulin pump | A bolus dose was less than X minutes ago | OpenAPS believes that the last bolus dose was more than X minutes ago
OpenAPS believes that it did not provide a dosing recommendation